MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.1

| Priority | 1 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Ultrasound (OBGY) | Installation Location | Obstetrics and Gynecology Ultrasound Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The system shall provide high‑resolution real‑time diagnostic ultrasound imaging for obstetrics, gynecology, and general abdominal examinations. - Equivalent or higher performance systems are acceptable. Brand‑neutral and performance‑based specifications shall apply. |  |  |  |  |  |
| - General Requirements [Must] |  |  |  |  |  |
| - Cart-based diagnostic ultrasound system suitable for continuous clinical use - Real-time 2D (B-mode) imaging with M-mode - Color Doppler, Power Doppler, and Pulsed-Wave (PW) Doppler - High-resolution color monitor and intuitive operator interface - Integrated image storage/archiving and reporting workflow - Stable performance suitable for high patient volume OPD use - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Technical Specifications [Must] |  |  |  |  |  |
| - Imaging Performance – B-Mode / M-Mode [MUST] - High-resolution grayscale imaging optimized for abdominal and pelvic examinations - Tissue Harmonic Imaging (THI) or equivalent - Speckle reduction / noise reduction imaging or equivalent - Spatial compounding (multi-angle) or equivalent - Dynamic range adjustment and multi-frequency capability per probe - Depth: suitable for obstetric abdominal scanning (deep penetration) while maintaining image clarity - Freeze, zoom, cine loop review with frame-by-frame control - M-mode for fetal heart assessment and motion evaluation - Doppler Performance [MUST] - Color Doppler with adjustable PRF, gain, wall filter, and color box size/position - Power Doppler for low-flow detection - PW Doppler with adjustable sample volume and angle correction - Duplex/Triplex capability (B + Color + PW) as supported by system - Doppler measurement packages for OB/GYN (e.g., uterine artery, umbilical artery as applicable) - Clear spectral display with velocity scale and baseline controls - OB/GYN Measurement & Clinical Packages [MUST] - The system shall provide an integrated OB/GYN measurement package including, at minimum: - Obstetric biometry: BPD, HC, AC, FL and standard fetal growth calculations - Gestational age estimation, EDD calculation, fetal weight estimation (standard formulas supported by system) - Fetal heart rate measurement - Placenta and amniotic fluid assessment tools (e.g., AFI) - Gynecology measurements: uterus, endometrium, ovaries/adnexa measurements - Built-in annotation and body marks appropriate for OBGY - Display & User Interface [MUST] - High-resolution color monitor ≥ 19 inch - Adjustable monitor height/tilt/swivel - Control panel with ergonomic layout suitable for long scanning sessions - One-touch optimization keys (e.g., gain, depth, focus) or equivalent workflow keys - QWERTY keyboard or equivalent for patient data entry and reporting - Multi-language support (English mandatory) - System Performance & Workflow [MUST] - Fast system boot-up suitable for OPD workflow - Cine memory capability for review of dynamic sequences - Preset management for OB abdomen, OB transvaginal, GYN pelvis and Doppler presets - Patient database management with search and retrieval - On-screen help or guided workflow features acceptable - Transducers / Probes [MUST] System shall be supplied with probes suitable for OBGY use. At minimum: A) Convex abdominal probe - Frequency range approximately 2–6 MHz (or equivalent range) - Suitable for routine obstetric abdominal scanning and deep penetration B) Endocavitary (transvaginal) probe - Frequency range approximately 5–9 MHz (or equivalent) - Suitable for early pregnancy and gynecologic pelvic scanning C) Linear probe - Frequency range approximately 5–12 MHz (or equivalent) - Useful for superficial imaging as needed (e.g., breast/thyroid or superficial pelvic structures) - All probes must be compatible with the offered ultrasound system and support required modes (B/Color/PW) as applicable. - Connectivity & Data Management [MUST] - DICOM 3.0 capability for image export (Storage and Print or equivalent) - USB ports for data export - Network connectivity (LAN) for PACS/EMR integration as applicable - Capability to store still images and cine clips internally - Standard image formats export capability (e.g., JPG/AVI or equivalent) in addition to DICOM - Printing & Reporting [MUST] - Ability to generate and print OB/GYN reports - Compatibility with external printer - Measurement results and annotations included in printout/export - Power Requirements [MUST] - AC 100–240 V, 50/60 Hz - Proper grounding and surge protection requirement stated by bidder - System suitable for hospital electrical environment |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| Each system shall include all items required for immediate clinical use: - Ultrasound main unit with monitor and control panel – 1 - Abdominal convex probe – 1 - Transvaginal probe – 1 - Linear probe – 1 - Probe holders and heated gel bottle holder – 1 set - Footswitch (if provided by manufacturer) – 1 - Thermal/B&W printer with associated cables to connect printer with the main unit and power cord-1 set - PC with Reporting printer with cable and power cord-1 set - Power cable – 1 - User manual and service manual – 1 set - Starter consumables (ultrasound gel) – as per bidder’s standard supply - Warranty certificate and calibration/quality documents – 1 set |  |  |  |  |  |
| - Optional Features (Price separately) |  |  |  |  |  |
| The following options should NOT be included in the base price and must be quoted separately: - 3D/4D imaging package - Elastography package - Contrast‑Enhanced Ultrasound (CEUS) - Advanced microvascular or high‑sensitivity Doppler - Additional probes or specialized software modules |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by certified personnel - System supplied fully assembled and ready to operate - On‑site acceptance testing included - Electrical safety testing included - Verification of all imaging modes (B/M/Color/Power/PW Doppler) - Verification of probe functionality and image quality - DICOM/USB export test (sample export) - Acceptance testing checklist completed and signed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for doctors/sonographers and nurses - Training includes: scanning presets, Doppler use, measurements, reporting, data export, basic troubleshooting - Biomedical/engineering orientation on routine checks and safe handling |  |  |  |  |  |
| - Warranty & After‑Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Includes parts, labor, probes, software, and travel costs - Preventive maintenance included during warranty period - Local authorized service support in Nepal (or formal service arrangement) required - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Documentation [MUST] |  |  |  |  |  |
| - Operating manual (hard & soft copy) - Service manual - Training materials |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-37 (Particular requirements for diagnostic ultrasound equipment) - CE marking and FDA clearance or equivalent international regulatory approval Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/