MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.10

| Priority | 1 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | LDR Bed (Labor, Delivery, Recovery) | Installation Location | Delivery Room |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The LDR bed shall be used to support mothers during labor, delivery, and recovery within the labor room. - The system shall provide safe patient positioning, caregiver ergonomics, and reliable operation for obstetric procedures. - Equivalent or higher performance systems are acceptable. Specifications shall be vendor-neutral and performance-based. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electrically operated obstetric bed - Stable and reliable structure suitable for daily clinical use - Easy-to-clean and disinfectant-resistant surfaces - Smooth and quiet motorized movements - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Core Functions & Movements [MUST] |  |  |  |  |  |
| - Electric height adjustment - Electric backrest adjustment - Trendelenburg and reverse Trendelenburg positioning - Removable or retractable foot section for delivery access - Manual emergency/backup control in case of power failure - CPR quick-release function |  |  |  |  |  |
| - Dimensions & Capacity [MUST] |  |  |  |  |  |
| - Patient surface length ≥ 2000 mm - Patient surface width ≥ 800 mm - Height adjustment range approximately 450 – 900 mm (or wider) - Safe working load ≥ 200 kg |  |  |  |  |  |
| - Mattress & Patient Support [MUST] |  |  |  |  |  |
| - Pressure-relief or high-density foam mattress - Seamless, waterproof, disinfectant-resistant cover - Mattress sections compatible with removable foot section |  |  |  |  |  |
| - Safety & Mobility [MUST] |  |  |  |  |  |
| - Lockable castors with central braking system - Side rails or patient supports for safety - Stable base design to prevent tipping - Hand grips or supports for labor assistance |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz (universal) - Internal battery backup or manual operation capability |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| All components required for full clinical operation must be included. - Each unit shall include: - Main LDR bed unit - Mattress set - Removable foot section - Adjustable leg supports (stirrups or calf supports) – 1 pair - IV pole – 1 - Fluid basin or collection tray – 1 - Hand grips/supports – 1 set - Power cable - All required mounting hardware - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| The following shall NOT be included in the base price and must be quoted separately: - Additional mattresses or pads - Additional leg supports or specialty accessories - Integrated scale or monitoring modules - Additional storage drawers or furniture accessories |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional and safety testing on-site - Device supplied fully assembled and ready for use - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for clinicians and nurses - Training on positioning, cleaning, and routine maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance during warranty included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |