MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.11

| Priority | 1 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Ultrasound (Labor Room) | Installation Location | Labor Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - This ultrasound system shall be used in the labor and delivery room for rapid obstetric, gynecologic, and basic abdominal diagnostic imaging. - The system shall provide reliable real-time imaging suitable for emergency and bedside use. - Equivalent or higher performance systems are acceptable. Specifications shall be vendor-neutral and performance-based. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade diagnostic ultrasound system suitable for bedside/point-of-care use - Cart-based compact or portable system with easy mobility - Real-time B-mode imaging with M-mode - Color Doppler and PW Doppler capability - Designed for rapid clinical decision-making - Durable housing resistant to frequent cleaning/disinfection - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Technical Specifications [Must] |  |  |  |  |  |
| - Portability & Mobility [MUST] - Compact footprint suitable for crowded labor rooms - Integrated trolley with lockable wheels - Lightweight and easy to move between beds - Stable base to prevent tipping - Cable management for safe transport - Quick relocation without recalibration - Boot-Up & Workflow Performance [MUST] - Fast system boot-up suitable for emergency use - Preset OB quick-exam mode available - One-touch optimization or automatic image optimization - Minimal steps from power-on to scanning - Rapid patient data entry - Imaging Performance – B-Mode / M-Mode [MUST] - High-resolution grayscale imaging optimized for obstetric use - Tissue Harmonic Imaging or equivalent - Speckle reduction / noise reduction processing - Spatial compounding or equivalent - Depth and focus adjustment suitable for abdominal OB imaging - Cine loop review with freeze and playback - M-mode for fetal cardiac activity assessment - Doppler Performance [MUST] - Color Doppler and Power Doppler capability - PW Doppler with adjustable sample volume - Adjustable PRF, gain, wall filter, and angle correction - Clear spectral waveform display - Duplex capability (B + Doppler) - OB Clinical Measurement Package [MUST] - System shall include standard obstetric measurement functions, including at minimum: - Fetal heart rate (FHR) - Basic fetal biometry (BPD, HC, AC, FL or equivalent) - Gestational age estimation - Placenta location documentation - Amniotic fluid assessment (e.g., AFI or equivalent) - Built-in annotation/body markers for OB use - Display & User Interface [MUST] - Color monitor ≥ 15 inches (compact systems acceptable) or ≥ 19 inches (cart type) - Adjustable screen angle/height - Ergonomic control panel - Backlit keyboard or equivalent - Dedicated shortcut keys for depth, gain, freeze, store - Multi-language support (English mandatory) - Transducers / Probes [MUST] At minimum: A) Convex abdominal probe - Frequency range approximately 2–6 MHz or equivalent - Suitable for obstetric abdominal examinations B) Endocavitary probe (transvaginal) - Frequency range approximately 5–9 MHz or equivalent - For early pregnancy and pelvic evaluation - All probes shall support required imaging and Doppler modes. C) Linear probe - Frequency range approximately 5–12 MHz (or equivalent) - Useful for superficial imaging as needed (e.g., breast/thyroid or superficial pelvic structures) - All probes must be compatible with the offered ultrasound system and support required modes (B/Color/PW) as applicable. - Data Storage & Connectivity [MUST] - Internal image and cine storage capability - USB ports for export - LAN connectivity for PACS/EMR (DICOM compatible) - DICOM 3.0 image export capability - Standard image/video export formats supported - Power System & Battery [MUST] - AC 100–240 V operation - Internal rechargeable battery - Battery runtime ≥ 60 minutes continuous scanning - Battery level indicator - Automatic safe shutdown protection - Suitable for short transport or power interruption during emergency - Infection Control & Durability [MUST] - Smooth, easy-to-clean surfaces - Resistant to standard hospital disinfectants - Sealed keyboard or wipeable interface - Robust construction for high-use clinical areas |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| Each unit shall include: - Ultrasound main unit with trolley – 1 - Convex probe – 1 - Endovaginal probe – 1 - Linear Probe-1 - Probe holders – 1 set - Heated Gel bottle holder – 1 - Power cable – 1 - User manual and service manual – 1 set - Starter consumables – as per bidder standard supply |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - On-site installation and commissioning - Verification of imaging quality and probe functionality - Doppler performance verification - Battery runtime test - DICOM/USB export test - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for OB doctors and nurses - Training includes rapid exam workflow, measurements, Doppler use, and data export - Basic maintenance and cleaning instruction |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance included - Response time ≤ 48 hours - Spare parts availability ≥ 7 years - Local or officially supported service arrangement required |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-37 (Diagnostic Ultrasound Safety) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/