MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.12

| Priority | 1 | 2 | 3 | Floor | GF,1F-DR, 2F, 3F | Quantity | 1EA | 2EA | 6EA |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  |  | NIBP Monitor | Installation Location | Labor Room |  |  |  |
|  |  |  |  |  |  | Surgery Examination Room Neurosurgery Treatment Room |  |  |  |
|  |  |  |  |  |  | Treatment Room (2F) |  |  |  |
|  |  |  |  |  |  | Treatment Room (3F) |  |  |  |
| Specification Level |  |  |  | Mid | Remarks |  |  |  |  |
| General Specification |  |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |  |  |
| - The non-invasive blood pressure monitor shall be used for routine and intermittent measurement of arterial blood pressure and pulse rate in adult, pediatric, and neonatal patients. - The device shall provide oscillometric blood pressure measurements suitable for routine clinical use in wards, outpatient departments, and emergency rooms, and shall be designed for frequent intermittent monitoring in ward and outpatient environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |  |  |
| - Medical-grade non-invasive blood pressure monitor (standalone type) - Oscillometric measurement principle - Suitable for adult, pediatric, and neonatal patients - Compact and portable design - Stable operation for repeated clinical use - Easy-to-clean surfaces resistant to hospital disinfectants - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |  |  |  |  |
| - Measurement method: oscillometric - Blood pressure measurement parameters: - Systolic blood pressure (SBP) - Diastolic blood pressure (DBP) - Mean arterial pressure (MAP) - Measurement range suitable for adult, pediatric, and neonatal patients - Measurement accuracy: ±3 mmHg or better (or ±2% of reading, whichever is greater) - Automatic measurement and manual measurement modes selectable - Adjustable automatic measurement intervals - Reliable performance during patient movement within acceptable limits |  |  |  |  |  |  |  |  |  |
| - Patient Modes [MUST] |  |  |  |  |  |  |  |  |  |
| - Adult mode - Pediatric mode - Neonatal mode - Automatic cuff pressure control according to selected patient mode - Over-pressure protection for neonatal use |  |  |  |  |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |  |  |  |  |
| - Clear digital display showing SBP, DBP, MAP, and pulse rate - Display size suitable for bedside viewing - Visual indicators for measurement status - Simple keypad or touch interface - Error messages displayed for motion, cuff leak, or abnormal measurement |  |  |  |  |  |  |  |  |  |
| - Alarm Functions [MUST] |  |  |  |  |  |  |  |  |  |
| - High and low blood pressure alarm limits adjustable - Pulse rate alarm - Alarm activation during abnormal measurement conditions - Audible and visual alarm indicators - Alarm silence function with automatic reactivation |  |  |  |  |  |  |  |  |  |
| - Memory & Trend [MUST] |  |  |  |  |  |  |  |  |  |
| - Internal memory storage for measurement records (minimum 200 readings) - Time and date stamping of stored measurements - Review of historical readings on device display - Trend review capability to observe BP changes over time |  |  |  |  |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz operation via adapter - Internal rechargeable battery - Battery runtime ≥ 6 hours under typical intermittent measurement use - Battery charging indicator - Low battery alarm |  |  |  |  |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |  |  |  |  |
| - Automatic over-pressure protection - Cuff leak detection - Motion artifact detection - Automatic measurement termination under unsafe conditions - Defibrillation-protected patient connection (if applicable) |  |  |  |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |  |  |  |
| - Each unit shall include all components required for immediate clinical use: - NIBP monitor main unit – 1 - Adult cuff – 1 - Large adult cuff – 1 - Pediatric cuff – 1 - Neonatal cuff – 1 - Rechargeable battery (installed) - AC power adapter – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |  |  |  |  |
| - On-site functional testing - Verification of measurement accuracy and alarm functions - Acceptance testing checklist completed |  |  |  |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |  |  |  |
| - On-site user training for nurses and clinical staff - Training includes cuff selection, measurement modes, alarm setting, cleaning, and basic troubleshooting |  |  |  |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |  |  |  |
| - Comprehensive warranty minimum 2 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance during warranty period - Response time ≤ 48 hours - Spare parts availability ≥ 5–7 years |  |  |  |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic Compatibility) - IEC 80601-2-30 (Particular requirements for NIBP monitors) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |  |  |

/End/