MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.13

| Priority | 1 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Neonatal Bubble CPAP System | Installation Location | NICU |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - This equipment is intended for neonatal and pediatric respiratory support, delivering non‑invasive Bubble CPAP therapy using heated, humidified blended gas. - It shall provide accurate control of FiO₂, stable CPAP delivery, and safe humidification for NICU and special newborn care units. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Low flow blender with mixed gas (FiO2 21-100%) output upto ≥ 11 LPM that incorporates cascaded flowmeters (1 LPM + 10 LPM) for precise low-flow adjustment. - Bubble CPAP generator capable of delivering continuous positive airway pressure (CPAP) through an underwater column system compatible with nasal prongs, masks and neonatal circuits. - Heated humidification system with non-invasive mode. |  |  |  |  |  |
| - Ventilation Modes & Performance [MUST] |  |  |  |  |  |
| - Air-Oxygen Blender: - FiO2 range: 21-100% adjustable - Total Flow Capacity: ≥ 11 L/min - Flowmeter range: 1LPM & 10LPM Cascaded - Mixing Method: Pneumatic proportioning - Input Gas: Input gas pressure compatible with standard hospital medical gas pipeline systems (typically 3–5 bar). - Bubble CPAP Generator: - CPAP Levels: 3-12 cmH2O - Input Flow: 4-15L/min - Auto-leveling bubble chamber - Water Column Depth: Marked increments; stable bubbling - Pressure-relief manifold - Compatible with general heated humidification systems - Heated Humidifier - Modes: Invasive/Noninvasive - Chamber temperature control: Dual sensor servo-controlled - Humidity output: Minimum 33 mg/L in NIV mode |  |  |  |  |  |
| - Display, Alarms & Safety [MUST] |  |  |  |  |  |
| - LCD ≥ 5 inches - Waveform display (pressure/flow) and numeric parameters - Audible and visual alarms including: - High/low pressure - Low minute ventilation (or apnea) - Disconnection/leak - Power failure / low battery - High respiratory rate (or equivalent) - Alarm limits adjustable |  |  |  |  |  |
| - Humidification [MUST] |  |  |  |  |  |
| - Compatible with heated humidifier system - Humidifier (integrated or external) shall be included |  |  |  |  |  |
| - Power & Portability [MUST] |  |  |  |  |  |
| - AC power: 100–240V, 50/60Hz - Internal rechargeable battery backup ≥ 4 hours (minimum) - Carry handle or trolley-compatible design - Automatic battery charging when connected to AC power |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - All components required for full clinical operation must be included. - Each unit shall include: - Air oxygen blender – 1 unit - Bubble CPAP generator kit – 5 units - Heater Humidifier unit- 1 unit - Reusable humidifier chamber – 1 unit - Heater wire and temperature sensor – 1 unit extra - User manual and service manual (both hardcopy and soft copy) mandatory for each main unit |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional circuits and consumables - Trolley/cart - Data export / connectivity module (if offered) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional testing on-site - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for clinicians and nurses - Training shall cover modes, alarm management, cleaning, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/