MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.14

| Priority | 1 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Ventilator (HFO+CMV) | Installation Location | NICU |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The ventilator shall provide invasive ventilation for neonatal and pediatric patients, and where applicable adult patients, in critical care environments, including both conventional ventilation modes and High Frequency Oscillatory Ventilation (HFO) with volume guarantee. - The system shall support precise low tidal volume ventilation suitable for neonatal applications and shall be designed for continuous clinical operation in hospital environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - ICU/NICU grade ventilator suitable for continuous operation - Supports neonatal and pediatric ventilation (adult capability acceptable) - Integrated display with waveforms, loops, and alarms - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Ventilation Modes & Performance [MUST] |  |  |  |  |  |
| - Conventional Modes - Volume control ventilation (VCV) and pressure control ventilation (PCV) - SIMV (volume or pressure) + PSV (or equivalent) - CPAP/PSV (for spontaneous breathing support) - Assist Control - HFO Mode - High Frequency Oscillatory Ventilation (HFO) mode (both Invasive and Non-invasive) with volume guarantee included - Adjustable frequency range suitable for neonatal use - Adjustable amplitude and mean airway pressure control - Neonatal Performance - Neonatal tidal volume range suitable for very low birth weight patients (e.g., VT down to ≤ 2 mL) - FiO2 21–100% - Flow/pressure sensors suitable for neonatal circuits - Automatic leak compensation / trigger sensitivity suitable for neonates (where applicable) |  |  |  |  |  |
| - Monitoring & Graphics [MUST] |  |  |  |  |  |
| - Color display ≥ 12 inches - Real-time waveforms (pressure/flow/volume) - Loops (pressure-volume, flow-volume) or equivalent - Trending of key parameters ≥ 72 hours (or equivalent internal memory) - Event log and alarm log |  |  |  |  |  |
| - Alarms & Safety [MUST] |  |  |  |  |  |
| - Audible/visual alarms including high/low pressure, apnea, minute ventilation, oxygen, power failure - Adjustable alarm limits - Safety valve / overpressure protection suitable for neonatal ventilation - Backup ventilation function (if primary mode fails) |  |  |  |  |  |
| - Power & Gas Supply [MUST] |  |  |  |  |  |
| - AC power 100–240V, 50/60Hz - Battery backup ≥ 2 hours minimum - Operates with oxygen and air supply (pipeline) and/or internal turbine (either acceptable) - Compatible with standard hospital gas connectors (to be confirmed during installation) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| All components required for full clinical operation must be included. - Each unit shall include: - Main ventilator unit - Humidification system (heated humidifier) with chamber – 1 set - Neonatal circuit set – minimum 2 sets - Pediatric circuit set – minimum 2 sets - HFOV neonatal circuit-10 - Flow sensors appropriate for neonatal use with adapters and wires (if required)-minimum 2 - Oxygen sensor (if required by the system) - Power cable - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Adult circuit sets - Advanced monitoring modules (e.g., integrated CO2) if offered - Transport trolley/cart (if not included) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, commissioning by qualified personnel - On-site functional testing and alarm verification - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for clinicians and biomedical staff - Training includes mode setup (conventional + HFO), alarm management, cleaning, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/