MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.15

| Priority | 1 | Floor | 1F-OT | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Operating Light | Installation Location | Operating Room 1,2,3 |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The operating light shall provide bright, uniform, and shadow-less illumination suitable for surgical procedures in the operating room environment. - The light field shall provide stable during repositioning without noticeable intensity fluctuation and shall support reliable operation during routine surgical use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Ceiling mounted LED surgical light system - Diameter of Dome: 700 mm(main) and 500mm(satellite) - Sterilizable handle for positioning - Smooth multi-axis positioning and stable arm - Each Light head is controlled by a separate control panel with waterproof film touch technology for operator convenience and easy cleaning and disinfection. - Disinfectant-resistant surfaces - Quantity of bulbs: at least 160 in main dome and at least 80 LED bulbs - Temperature rise in surgeon’s head: 1℃ - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Light Performance [MUST] |  |  |  |  |  |
| - Central illuminance ≥ 160,000 lux (at nominal working distance) - Color rendering index (CRI) ≥ 97 - Adjustable light intensity - Color temperature 10 level adjustable, adjustment range 3500-5800K or wider; - Low heat emission at surgical field - Shadow less technology - There are 5 independent green auxiliary lighting sets for endoscopic surgical lighting with illumination not exceeding 500lx. - Uniform light field suitable for general surgery and OB/GYN procedures |  |  |  |  |  |
| - Mechanical & Control [MUST] |  |  |  |  |  |
| - Multi-joint arm with wide positioning range - Stable positioning without drift - Control via wall panel/handle (either acceptable) |  |  |  |  |  |
| - Electrical & Safety [MUST] |  |  |  |  |  |
| - AC power suitable for hospital supply (100–240V, 50/60Hz) - Electrical safety compliant with IEC 60601-1 or equivalent - Each LED lighting group module has an independent circuit control. When any group fails, the Surgical Light should recover its illumination within 5s to ensure safe use during the operation. - LED lifetime ≥ 65,000 hours |  |  |  |  |  |
| - Configuration & Accessories [MUST] |  |  |  |  |  |
| All necessary mounting hardware and accessories required for installation must be included. - Dual-head configuration (main + satellite) OR equivalent configuration suitable for OR use - Ceiling arms and mounting hardware - Sterilizable handles - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Integrated camera/recording module - Additional sterile handle sets - Additional monitor/arm accessories |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Site survey (as required), delivery, installation, and commissioning by qualified personnel - Functional testing on-site - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for OR staff on safe use and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-41 (Particular requirements for surgical luminaires and luminaires for diagnosis) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/