MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.16

| Priority | 1 | Floor | 1F-OT | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Anesthesia Machine | Installation Location | Operating Room 1,2,3 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The anesthesia workstation shall be used for delivery of inhalational anesthesia and integrated mechanical ventilation for adult and pediatric patients during surgical procedures. - The system shall support low-flow anesthesia operation and provide stable performance suitable for operating room environments and shall be designed for continuous clinical operation in hospital environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Integrated anesthesia machine with ventilator and monitoring interface - Supports adult and pediatric patients (neonatal capability acceptable) - User-friendly interface suitable for OR workflow - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Gas Delivery & Safety [MUST] |  |  |  |  |  |
| - Electronic or mechanical gas mixing system (either acceptable) with minimum O2 and Air capability - FiO2 control with hypoxic guard / minimum O2 ratio protection (or equivalent safety mechanism) - O2 flush function - Pipeline gas connection for O2 and Air (and N2O if used in facility; optional) - Built-in safety alarms for gas supply failure / low O2 |  |  |  |  |  |
| - Vaporization [MUST] |  |  |  |  |  |
| - Minimum two vaporizer mounting positions (Selectatec type or equivalent) - Vaporizers shall be compatible with commonly used agents (Isoflurane and Sevoflurane) |  |  |  |  |  |
| - Ventilation (Integrated Ventilator) [MUST] |  |  |  |  |  |
| - Ventilation modes suitable for OR: - Volume control (VCV) - Pressure control (PCV) - SIMV (volume or pressure) and PSV (or equivalent) - Tidal volume range suitable for adult and pediatric use - PEEP capability - Adjustable respiratory rate and inspiratory time - Integrated alarms: apnea, high/low pressure, volume, oxygen, power failure |  |  |  |  |  |
| - Monitoring & Display [MUST] |  |  |  |  |  |
| - Color display ≥ 10 inches (or equivalent) - Display of key ventilation parameters and waveforms - Integrated or compatible monitoring for: - Side stream EtCO2 (capnography) OR capability to add EtCO2 module - Agent monitoring for 5 agents |  |  |  |  |  |
| - Power & Backup [MUST] |  |  |  |  |  |
| - AC 100–240V, 50/60Hz - Internal battery backup ≥ 120 minutes minimum (or equivalent) - Automatic charging |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| All components required for safe clinical operation must be included. - Main anesthesia workstation with ventilator - Universal Vaporizer Manifold - Isoflurane Vaporizer - Sevoflurane Vaporizer - AGM module - Breathing circuits (adult and pediatric) – minimum 2 sets each - Spare parts: 5× flow sensors, 1× bellows, free O2 sensor replacement during warranty - Sensors and disposables – starter kit - CO2 sampling line - Flow sensors / breathing bag / hoses (as required) - Power cable - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional circuits and consumables - Scavenging accessories (if not included as standard) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - On-site functional testing and safety checks - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for anesthetists, nurses, and biomedical staff - Training includes routine checks, alarm management, and basic troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-13 (Particular requirements for anesthesia workstations) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/