MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.17

| Priority | 1 | Floor | 1F-OT | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Patient Monitor (OT) | Installation Location | Operating Room 1, 2, 3 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The patient monitor shall provide continuous real-time monitoring of vital physiological parameters during anesthesia, surgical procedures, and critical care management for adult, pediatric, and neonatal patients. - The system shall be suitable for operating room and critical care environments and shall be designed for continuous clinical operation in hospital environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Bedside multi-parameter monitors suitable for continuous ICU/OR use - User-friendly interface with quick parameter setup - Most be Modular - Bispectral Index required - Durable construction and disinfectant-resistant housing - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Monitoring Parameters [MUST] |  |  |  |  |  |
| - ECG (3/5 lead or equivalent) - SpO₂ - Non-invasive blood pressure (NIBP) - Temperature (minimum 1 channel) - Invasive blood pressure (minimum 2 channels) - EtCO₂ (side stream) - BIS All above parameters shall be included in the base configuration. |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |
| - Color display ≥ 15 inches - Simultaneous waveform and numeric display - Minimum 6 waveforms displayed concurrently - Rotary knob control with touch screen - Clear alarm indicators and parameter labels |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |
| - ECG with heart rate and arrhythmia detection - SpO₂ with motion artifact resistance - NIBP oscillometric method with adult/pediatric/neonatal modes - IBP with standard transducer compatibility - Capnography with numeric and waveform display |  |  |  |  |  |
| - Alarms & Safety [MUST] |  |  |  |  |  |
| - Audible and visual alarms for all parameters - Adjustable alarm limits - Alarm priority differentiation - Battery and system fault alarms |  |  |  |  |  |
| - Data & Connectivity [MUST] |  |  |  |  |  |
| - Trend memory ≥ 72 hours (minimum) - Event recording - USB or LAN connectivity - Capability to connect to central monitoring station |  |  |  |  |  |
| - Power & Portability [MUST] |  |  |  |  |  |
| - AC 100–240V, 50/60Hz - Internal rechargeable battery ≥ 4 hours - Portable or trolley-compatible design |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - ECG cable - SpO₂ probe (reusable) - NIBP cuffs (adult, pediatric, neonatal) - IBP cables (minimum 2) - Temperature probe - EtCO₂ sampling lines starter pack - BIS starter pack - Power cable - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Central monitoring station - Additional parameter modules - Additional probes |  |  |  |  |  |
| - Warranty, Installation, Training |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Response time ≤ 48 hours - Spare parts availability ≥ 7 years - Delivery, installation, and commissioning by qualified personnel - On-site functional testing and safety checks - Acceptance testing included - The supplier must conduct user training on operational functions, routine checks, and maintenance |  |  |  |  |  |
| - Compliance |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/