MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.18

| Priority | 1 | Floor | 1F-OT | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Electro Surgical Unit (OT) | Installation Location | Operating Room 1, 2, 3 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The electrosurgical unit shall be used in the operating room for cutting and coagulation of tissue during surgical procedures, including general surgery and obstetrics/gynecology applications. - The system shall provide monopolar and bipolar electrosurgery with stable output performance, safe operation, and reliable energy delivery across varying tissue impedances, and shall be suitable for continuous surgical use in hospital operating room environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled electrosurgical generator for OR use - Monopolar and bipolar operation - Clear user interface with power settings and mode selection - Stable output with automatic control according to tissue impedance (or equivalent) - Disinfectant-resistant housing and easy cleaning design - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Operating Modes & Performance [MUST] |  |  |  |  |  |
| - Monopolar - Pure cut mode - Blend cut mode (or equivalent) - Coagulation modes: spray/fulguration and desiccation (or equivalent) - Output power suitable for general surgery (monopolar cut and coag each capable of high output, e.g., up to ≥ 300 W) - Bipolar - Bipolar coagulation mode - Output power suitable for routine bipolar procedures (e.g., up to ≥ 50 W) - Integrated Vessel Sealing module with 350W or equivalent vessel sealing power - Safety & Output Control - Automatic tissue sensing/impedance monitoring or equivalent output stabilization - Overload and short-circuit protection - Return electrode monitoring (REM) / patient plate contact quality monitoring or equivalent safety function |  |  |  |  |  |
| - User Interface & Alarms [MUST] |  |  |  |  |  |
| - Digital display of selected mode and output power - Audible and visual alarms including: - Patient plate/return electrode alarm - System fault alarm - Over-temperature/overload alarm - Volume control for alarms |  |  |  |  |  |
| - Footswitch & Activation [MUST] |  |  |  |  |  |
| - Footswitch control for monopolar cut/coag - Footswitch control for bipolar - Handpiece activation capability (if supported by design) - Footswitch shall be sealed and suitable for OR cleaning |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - The ESU supply shall include all components required for full clinical operation, minimum: - ESU main unit – 1 - Footswitch (cut/coag) – 1 - Bipolar Footswitch-1 - Monopolar pencil handpiece (reusable) – minimum 2 - Monopolar pencil handpiece (disposable) – starter pack (minimum 10) - Patient return electrodes/plates – starter pack (adult minimum 20) - Bipolar forceps – 1 - Bipolar cable – 1 - Monopolar cable set – 1 - Electrode tips set (assorted) – starter pack - Power cable – 1 - User manual and service manual – 1 set All required connectors/adapters must be included. “Accessories set” wording alone is not acceptable; quantities shall be supplied as above or better. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| The following shall NOT be included in the base price and must be quoted separately: - Argon plasma coagulation module (if offered) - Smoke evacuation unit (if offered) - Trolley/cart (if not included) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - On-site functional testing of all modes and alarms - Electrical safety testing on-site - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for surgeons, nurses, and technicians - Training shall include safe operation, patient plate placement, alarm management, cleaning, and basic troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance at least once per year during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-2-2 (Particular requirements for high-frequency surgical equipment) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/