MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.19

| Priority | 1 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Sterilizer Steam | Installation Location | Sterilization Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The steam sterilizer shall be used for sterilization of surgical instruments, linen packs, and reusable medical devices using saturated steam sterilization processes. - The system shall support safe and efficient CSSD workflow, including operation compatible with physical separation between contaminated and sterile areas, and shall be hospital-grade equipment suitable for continuous daily operation in hospital environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Configuration & Type [MUST] |  |  |  |  |  |
| - Horizontal steam sterilizer - Pre-vacuum (dynamic air removal) type - Double door (pass-through) configuration for clean/dirty area separation - Built-in electric steam generator shall be provided as part of the sterilizer system OR an external steam supply package shall be provided by the supplier. In either case, the supplier shall be fully responsible for providing a complete, functional steam supply system including all accessories, piping interfaces, and safety components required for EN 285-compliant operation. - Fully automatic cycle control - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Chamber Capacity & Structure [MUST] |  |  |  |  |  |
| - Effective chamber volume: minimum 250 liters - Stainless steel (316L) sterilization chamber, jacket and pipeline - Double-door pass-through design: - One door on contaminated side - One door on clean side - Interlocking mechanism preventing simultaneous opening - Hinged door mandatory - Chamber dimensions suitable for standard instrument trays and packs - Loading shelves/racks included (must be made up of SS316) - Trays/baskets included (minimum 3) NOTE: Capacity less than 250 liters or single-door configuration is NOT acceptable. |  |  |  |  |  |
| - Sterilization Performance & Cycles [MUST] |  |  |  |  |  |
| - Standard sterilization temperatures: 121°C and 134°C - Pre-vacuum sterilization cycles with air removal - Post-vacuum drying cycle - Adjustable sterilization time and temperature (within safe validated limits) - Bowie-Dick test cycle for vacuum performance verification - Leak test cycle - Fast cycle capability suitable for routine surgical instrument turnover |  |  |  |  |  |
| - Control System & Monitoring [MUST] |  |  |  |  |  |
| - Microprocessor/PLC-based automatic control - Digital display for: - Time - Temperature - Pressure - Cycle status - Automatic cycle recording via: - Built-in printer and/or - USB data export - Alarm history and cycle history storage - User access control to prevent unauthorized parameter changes |  |  |  |  |  |
| - Utilities & Installation Requirements [MUST] |  |  |  |  |  |
| - Electrical supply compatible with hospital power (bidder to specify rating) - Water supply connection: - The hospital shall provide a main water inlet point near the installation area. - The supplier shall provide all downstream connections, integration, and interfaces required for operation of the sterilizer and water treatment system. - Drainage connection included - All necessary hoses, connectors, valves, piping interfaces, and fittings required for complete installation shall be supplied by the bidder. - Noise and heat emissions suitable for indoor CSSD environment - Water treatment system: - The supplier shall provide and install a dedicated water treatment unit suitable for steam sterilizer operation. - The water treatment system shall include reverse osmosis (RO) or equivalent to meet the sterilizer manufacturer’s feed water quality requirements and applicable guidance under EN 285. - The system shall include all required components such as pre-filters, RO module, storage tank (if required), booster/transfer pump, pressure regulator, conductivity/TDS monitoring (or equivalent indicator), hoses/pipes, valves, and fittings, and shall be fully integrated and commissioned for stable daily operation. - The supplier shall ensure compatibility between treated water quality and sterilizer performance requirements. • Steam supply integration: - The supplier shall provide and install all necessary interfaces and components required for connection between the sterilizer and the steam generation/supply system (built-in or external as offered), ensuring stable operation compliant with EN 285 sterilization cycles. • Site Dimension Constraint: - Due to spatial limitations of the installation area and required CSSD workflow configuration, the sterilizer shall be a double-door (pass-through) type, and the overall external equipment width shall not exceed 1,350 mm. - The supplier shall ensure that the RO unit and any auxiliary steam components can be installed within the allocated plant/utility area without obstructing CSSD workflow and maintenance access. - The supplier shall verify site dimensions prior to manufacturing and delivery, and the offered equipment must fully comply with this dimensional requirement. • Commissioning requirement: - Installation and commissioning shall include verification of water treatment operation and successful sterilization cycle testing in accordance with manufacturer procedures and EN 285 requirements. |  |  |  |  |  |
| - Steam Supply System (Steam Generator) [MUST] |  |  |  |  |  |
| - The sterilizer shall be supplied with a complete steam generation/supply solution suitable for continuous hospital CSSD operation. - Option A (Built-in steam generator): - Integrated electric steam generator suitable for EN 285-compliant steam sterilization cycles. - Minimum 21 KW heater with 3 units of 7KW heaters - Automatic water level control, low-water protection, and over-temperature protection. - Easy access for routine maintenance and descaling; drain connection included. - Option B (External steam supply package): - If external steam is proposed, the supplier shall provide a complete package including steam control valve, pressure reducing/conditioning components, steam trap, separator/filter (as applicable), gauges/sensors, and all interfaces required for stable sterilizer operation. - Steam quality and operating conditions shall be compliant with EN 285 and the sterilizer manufacturer’s requirements. - The bidder shall clearly specify which option is offered and provide a schematic diagram of the steam supply configuration. |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Door interlock system preventing opening when chamber is pressurized - Mechanical and electronic over-pressure protection - Over-temperature protection - Low-water alarm - Power failure alarm - Emergency stop or safe shutdown function - Automatic safe venting before door unlocks |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - The sterilizer supply shall include, minimum - Main sterilizer unit – 1 - Double pass-through doors with interlock – integrated - Loading rack/shelves – 1 set - Sterilization trays/baskets – minimum 3 - Printer or data logger – 1 - Starter pack of printer paper/ink (if applicable) - Water inlet filters – 1 set - Door gaskets – 1 - Pneumatic valve for jacket, chamber and exhaust- – 1 set - Required installation materials (hoses, connectors, clamps, fittings) - Power cable or electrical connection kit - User manual and service manual – 1 set - Steam generator / steam supply package (built-in or external as offered) – 1 set (complete with all safety and control components) - Water treatment unit (RO or equivalent) – 1 set (complete package) - Water quality monitoring device (conductivity/TDS indicator or equivalent) – 1 set - All interconnection kits for steam/water/drain (hoses, pipes, valves, clamps, fittings) – 1 lot All items necessary for complete installation and operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional trays/baskets - Loading trolley/cart - Additional water treatment beyond supplied RO - UPS (if requested by hospital) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by certified personnel - Site preparation guidance provided before installation - On-site functional testing of all cycles - Validation support during commissioning (temperature/pressure printout verification) - Acceptance testing included - Commissioning shall include verification of RO system operation and delivered water quality indicators (as applicable). - Commissioning shall include verification of steam supply performance and successful completion of Bowie-Dick and leak test cycles with printout/data record. |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for CSSD staff and technicians - Training shall include: - Cycle selection - Proper loading/unloading - Routine checks - Cleaning and maintenance - Safety procedures - Basic troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance at least once per year during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - EN 285 (Large steam sterilizers) - IEC 61010-1 (Safety requirements for electrical equipment for measurement, control, and laboratory use) - CE marking or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/