MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.2

| Priority | 1 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Operating Table (ER) | Installation Location | ER Treatment Room |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - This operating table shall be used in the Emergency Room (ER) treatment area to support minor surgical procedures, emergency interventions, and patient positioning. - The system shall be robust, reliable, and suitable for continuous hospital use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electro-hydraulic or equivalent powered operating table - Stable and reliable operation suitable for emergency use - Smooth and quiet movements - Easy-to-clean and disinfectant-resistant surfaces - Compatibility with standard surgical accessories - Mobile table with central braking system - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Technical Specifications [MUST] |  |  |  |  |  |
| - Safety & Compliance - Electrical safety compliant with IEC 60601-1 or equivalent - Mechanical safety compliant with applicable international standards - Emergency stop and manual override functions - Audible or visual indicators for system status where applicable - Dimensions & Capacity - Tabletop width ≥ 500 mm - Tabletop length ≥ 2000 mm - Height adjustment range approximately 600 – 1000 mm (or wider) - Safe working load ≥ 180 kg in all positions - Movement & Positioning - Height adjustment (electric) - Trendelenburg and reverse Trendelenburg positioning - Lateral tilt (left/right) - Back, leg, and head section adjustment - Longitudinal sliding function ≥ 280 mm for C-arm access - Materials & Construction - Tabletop sections radiolucent where applicable - Column and base with corrosion-resistant metal housing - Mattress: removable, seamless, pressure-relief type |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| Accessories shall be fully compatible and supplied as a complete operational set. - Operating table main unit - Mattress set (pressure-relief, anti-static, waterproof or equivalent) - Central braking system with lockable castors - Wired or wireless hand control - Standard accessory rails - Power cable - Battery Backup: At least 1 hour of full operation on battery (or provision for emergency manual use) - Orthopedic set - All necessary mounting hardware and connectors - User and service manuals |  |  |  |  |  |
| - Optional Accessories (price separately) |  |  |  |  |  |
| The following items shall not be included in the base price and must be quoted separately: - Additional specialized surgical accessories (e.g., Neurosurgical sets) - Additional arm supports or leg supports beyond the standard set - Additional mattresses or pads |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional testing and safety verification on-site - System supplied fully assembled and ready for clinical use - On-site acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for ER staff - Training shall cover safe operation, positioning functions, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Includes parts, labor, probes, software, and travel costs - Preventive maintenance included during warranty period - Local authorized service support in Nepal (or formal service arrangement) required - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Documentation [MUST] |  |  |  |  |  |
| - User manual (hard & soft copy) - Service manual |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-46 (Particular requirements for operating tables) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/