MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.20

| Priority | 1 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Laparoscopy System | Installation Location | Operating Room 2 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The laparoscopy system shall be used for minimally invasive surgery by providing high-quality visualization, stable illumination, and integrated image recording capabilities suitable for surgical applications. - The system shall be suitable for general surgery and obstetrics/gynecology laparoscopic procedures and shall support reliable operation in hospital operating room environments during continuous surgical use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. All imaging components shall be from the same manufacturer or fully certified for compatibility and system integration. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Complete laparoscopy tower system including camera, light source, monitor, and recording capability - Medical-grade components suitable for OR use - Stable, mobile tower/trolley for OR workflow - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Imaging System Performance [MUST] |  |  |  |  |  |
| - Camera System - Resolution: 4K with upgradability to 3D - High sensitivity with low noise image performance - White balance and image enhancement functions (manufacturer standard) - Monitor - Medical-grade monitor size ≥ 32 inches - Resolution compatible with camera output (4K) - Multiple inputs (HDMI/SDI/DVI acceptable) - Must have 12G-SDI, multiple inputs and outputs and support Picture-in-picture - Image Processing - Must have 4K resolution along with 3D upgradability - Real-time image display with minimal latency - Color reproduction suitable for surgery - Must include digital color spectrum modulation mode such as NIR/ICG fluoroscopy with both digital and optical filters |  |  |  |  |  |
| - Light Source & Illumination [MUST] |  |  |  |  |  |
| - LED light source (or equivalent high-efficiency light source) - Adjustable light intensity - Long life light engine (≥ 20,000 hours recommended) or equivalent - Compatible light cable included |  |  |  |  |  |
| - Recording & Data [MUST] |  |  |  |  |  |
| - Integrated recording system (built-in or separate recorder) - Recording to USB storage (or equivalent) - Still image capture function - Basic patient data labeling (if available) |  |  |  |  |  |
| - Scope & Optics [MUST] |  |  |  |  |  |
| - Laparoscope telescope: 10 mm diameter, 30° direction of view – 1 - Laparoscope telescope: 10 mm diameter, 0-90° variable angle of view – 1 - Hysteroscope telescope: 3 mm approx., 30° direction of view – 1 - Autoclavable scope (or equivalent sterilizable design) - Light cable compatible with the telescope and light source – 1 - Camera head compatible with scope coupler – 1 - Scope coupler/adapter set – 1 |  |  |  |  |  |
| - Insuffolator [MUST] |  |  |  |  |  |
| - Suitable for both adult as well as pediatric use - Flow rate: Adult up to 50L/min, Pediatric 15L/min - Pressure: Adult up to 30mmHg, Pediatric 15mmHg - Overpressure cutoff with automatic & electronically controlled gas refill - Should have the ability to provide heated gas to the patient - Touch control panel - Resuable tubing with heating element – 1 |  |  |  |  |  |
| - Instrument [MUST] |  |  |  |  |  |
| - Inner Sheath, size 4.3 mm, with working channel for semirigid 5 Fr. operating instruments - Outer Sheath, size 5 mm - Biopsy and Grasping Forceps, double action jaws, semirigid, length 40 cm, 5 Fr diameter - Scissors, blunt, single jaws, semirigid, length 40 cm, 5Fr diameter - Bipolar Vaporization Electrode, semirigid, needle electrode angled 90°, length 42 cm, 5 Fr diameter - Bipolar High Frequency Cord, length 300 cm - Veress needle, with spring loaded blunt inner cannula, diameter 2.1mm, length 13cm – Qty 1 - Pyramidal trocar set, 11mm – Qty 4 - Pyramidal trocar set, 6mm – Qty 4 - Sleeve reducer, 11/5mm – Qty 1 - Clip on reducer, 11/5mm – Qty 1 - Maryland dissecting forceps – Qty 1 - Babcock forcep – Qty 1 - Bowel Grasping forceps – Qty 1 - Atraumatic Grasping forceps – Qty 1 - Fenestrated curved Fundus Grasper – Qty 1 - Unipolar curved right angle dissecting forceps – Qty 1 - Traumatic Grasping forceps – Qty 1 - Reusable Scissors curved Unipolar – 10 - Traumatic Grasping forceps – Qty 1 - Disposable scissor inserts – Qty 1 - Traumatic Grasping forceps – Qty 1 - L hook electrode – Qty 1 - Spatula electrode – Qty 1 - Clip Applicator 300 – Qty 1 - Unipolar HF cable – Qty 1 - Bipolar Hf Cable – Qty 1 - Needle holder - Qty 1 - Suction irrigation tube, with lateral holes, 5mm, 36cm -Qty 1 - Claw forceps,10mm -Qty 1 - Bipolar Maryland forceps -Qty 1 |  |  |  |  |  |
| - Tower/Trolley & Accessories [MUST] |  |  |  |  |  |
| - Mobile tower/trolley with locking wheels - Shelves/holders for components - Cable management (Medical isolation multiport power strip) included - All required power cords, signal cables, connectors included - User manuals and service manuals for all main components All items required for a complete functional laparoscopy set must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - 4K upgrade (if offered) - 3D laparoscopy module (if offered) - Fluorescence imaging module (if offered) - Additional telescopes (0°, 45°) or smaller diameter scopes - Additional monitors |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - On-site functional testing of camera, light source, monitor, recording - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for surgeons and OR staff - Training includes setup, white balance, recording, cleaning, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance during warranty included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/