MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.21

| Priority | 1 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Operating Microscope | Installation Location | Operating Room 5 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The operating microscope shall be used in operating rooms for microsurgical procedures requiring stable magnified visualization and precise illumination, including ENT, plastic and reconstructive surgery, neurosurgery/spine (as applicable), and other precision surgical operations. - The system shall allow precise repositioning without loss of focus during surgery and shall support reliable operation for continuous surgical use in hospital operating room environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - High optical clarity microscope system suitable for microsurgery - Stable, vibration-free structure and floor stand - Smooth maneuverability with multi-axis arm - Motorized zoom and focus (or equivalent) - Ergonomic design suitable for surgeon comfort - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Optical System & Magnification [MUST] |  |  |  |  |  |
| - Variable magnification range suitable for microsurgical procedures. - Zoom system shall be motorized or provide equivalent smooth and precise magnification control suitable for microsurgery. - Focus system shall be motorized or provide equivalent precise focusing capability to allow stable image adjustment during surgery. - The system shall provide a long working distance suitable for surgical operating fields. The bidder shall specify the working distance, and the system shall support deep cavity access without frequent repositioning. - Optical system shall provide high contrast, high-resolution visualization with accurate color reproduction suitable for microsurgical applications. It must have integrated green and orange filters. |  |  |  |  |  |
| - Illumination [MUST] |  |  |  |  |  |
| - LED illumination with adjustable brightness (xenon acceptable if provided) - Uniform illumination across surgical field - Backup illumination feature or safe continuation method (e.g., spare light module or rapid replacement) to minimize downtime |  |  |  |  |  |
| - Viewing Tube & Eyepieces [MUST] |  |  |  |  |  |
| - Binocular viewing tube for main surgeon with ergonomic adjustment (tiltable/foldable acceptable) - Widefield eyepieces with diopter adjustment - Assistant observation tube (co-observation) included (side-by-side or face-to-face acceptable) |  |  |  |  |  |
| - Stand, Maneuverability & Controls [MUST] |  |  |  |  |  |
| - Stable floor stand with lockable wheels - Brakes for all axes (electromagnetic or mechanical acceptable) - Fine movement control via handgrips and footswitch (either acceptable) - Footswitch included for key functions (zoom/focus and/or light intensity) |  |  |  |  |  |
| - Video Output & Documentation [MUST] |  |  |  |  |  |
| - Integrated or attachable camera/video output for documentation and teaching - Digital video output interface (HDMI/SDI or equivalent) - If recording is not included in base, bidder shall quote recording module as Optional (price separately) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - The microscope supply shall include, minimum: - Microscope head – 1 - Floor stand – 1 - Main surgeon binocular tube – 1 - Assistant binocular tube – 1 - Eyepieces – 2 pairs (total 4 eyepieces) or equivalent configuration for main + assistant - Footswitch – 1 - Sterilizable knobs/caps – 1 set - Dust cover – 1 - Power cable and required connectors - User manual and service manual – 1 set All components required for full clinical operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| The following shall NOT be included in the base price and must be quoted separately: - 4K camera system upgrade (if offered) - Digital recording system - Monitor (≥ 22 inch) mounted on stand - Vacuum draping or drape air removal system (if offered) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - On-site functional testing of optics, illumination, brakes, and controls - Electrical safety testing on-site - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for surgeons and OR staff - Training includes positioning/balancing, use of footswitch, cleaning, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance during warranty included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/