MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 25

| Priority | 1 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Operating Table (OT 5) | Installation Location | Operating Room 5 |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The operating table shall be used in operating rooms to support neurological and ENT surgeries. The system shall provide stable positioning, imaging compatibility suitable for intraoperative imaging systems (e.g., C-arm fluoroscopy), and ergonomic access for surgical teams. - Table performance shall provide full functionality across the specified load range without restriction of movement or positioning functions. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electro-power hydraulic operating table - Floor-mounted mobile base with auto lock system - Stable T column design providing comfortable working space for surgeons - Smooth and quiet movements - Easy-to-clean and disinfectant-resistant surfaces - Stainless steel or corrosion-resistant construction - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Patient Weight Capacity [MUST] |  |  |  |  |  |
| - Static load capacity ≥ 350 kg - Articulated/operational load capacity ≥ 290 kg Equivalent or higher performance systems meeting the above functional and safety requirements are acceptable. |  |  |  |  |  |
| - Tabletop & Imaging Compatibility [MUST] |  |  |  |  |  |
| - X Ray access and radiolucent table top - C-arm compatible with unobstructed imaging access - Modular tabletop sections |  |  |  |  |  |
| - Positioning & Movements [MUST] |  |  |  |  |  |
| - Height adjustment range approximately 495–1000 mm or wider - Trendelenburg and reverse Trendelenburg ±30° minimum - Head plate and leg plates should be interchangeable. - Lateral tilt ±25° minimum - Headplate up 60° and down 90° - Back section up ≥ 85° and down ≥ 45° - Leg section up ≥ 20° and down ≥ 90° - Leg section split 180° - Motorized longitudinal slide ≥ 350 mm or more to both head and foot - Kidney bridge elevation ≥ 130 - Flex/Reflex 225°/95° |  |  |  |  |  |
| - Control & Safety Systems [MUST] |  |  |  |  |  |
| - Hand pendant control - Backup control panel on column or base - Battery backup for operation during power failure upto 85 operations - Battery charge indicator - Emergency stop or zero position button and manual override function |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Main operating table unit – 1 - Hand control – 1 - Integrated Side Rails-1 set - Backup control panel – integrated - Clamps/adapters – minimum 6 - Headrest (dual articulating) – 1 - Removable leg section-1 - Leg supports– 1 pair - IV pole – 1 - Safety straps – 1 set - Mattress set – pressure relief type - Power cable – 1 - User and service manuals – 1 set - Accessories for Neurosurgery procedures - OT Table adapter-1 - Adjustable base-1 - 3 pin headclamp-1 - Swivel adapter-1 - 3 reusable skull pins (pediatric and adult)- 1set - Horseshoe headrest with pads-1 set - All accessories required for full clinical use must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Pediatric accessories - Additional arm boards or supports |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional and load testing on-site - Electrical safety testing included - Acceptance testing included - Final positioning shall be coordinated with ceiling-mounted equipment to ensure full table movement without obstruction. |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for surgeons and OR staff - Positioning and safety procedures training |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA registration or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/