MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.26

| Priority | 1 | 3 | Floor | 1F-OT | Quantity | 1EA | 1EA |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  | Ultrasonic Surgical Energy Generator | Installation Location | Operation Room 2 |  |  |
|  |  |  |  |  | Operation Room 3 |  |  |
| Specification Level |  |  | High | Remarks |  |  |  |
| General Specification |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |
| - The ultrasonic surgical generator shall provide simultaneous cutting and coagulation using ultrasonic mechanical energy for open and minimally invasive surgical procedures. - The system shall enable precise tissue dissection with controlled thermal spread and reliable vessel sealing performance suitable for advanced surgical energy applications. - The equipment shall be suitable for laparoscopic and open surgical applications including general surgery, gynecology, and urology, and shall function as an advanced energy device supporting modern surgical workflow. - The ultrasonic surgical generator is intended to function as an advanced energy device complementary to conventional electrosurgical systems used in the operating room. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |
| - Ultrasonic energy-based surgical generator - Capable of simultaneous cutting and coagulation - Microprocessor-controlled output regulation - Compatible with reusable or disposable ultrasonic handpieces - Designed for continuous operating room use - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |
| - Energy Performance [MUST] |  |  |  |  |  |  |  |
| - Ultrasonic operating frequency within the typical ultrasonic surgical range of approximately 50–60 kHz or manufacturer equivalent technology. - Mechanical vibration-based tissue dissection - Reduced thermal spread compared to conventional electrosurgery - Stable energy delivery under variable tissue impedance - System shall allow controlled dissection with minimal smoke production |  |  |  |  |  |  |  |
| - Vessel Sealing Capability [MUST] |  |  |  |  |  |  |  |
| - Reliable sealing of vessels up to 5 mm diameter or greater as specified by the manufacturer. - Consistent sealing performance without carbonization - Automatic energy modulation based on tissue feedback |  |  |  |  |  |  |  |
| - Output Control & Modes [MUST] |  |  |  |  |  |  |  |
| - Multiple selectable power levels - Separate cutting and coagulation performance profiles - Automatic power adjustment during activation - Audible activation tone - Clear visual indication of active mode |  |  |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |  |  |
| - Automatic shutdown during overload or overheating - Handpiece recognition and compatibility control - System self-test at startup - Activation safety lock or standby mode - Electrical isolation compliant with medical safety standards |  |  |  |  |  |  |  |
| - User Interface [MUST] |  |  |  |  |  |  |  |
| - Digital display showing power level and system status - Simple control interface suitable for surgical workflow - Footswitch and/or hand activation compatibility - Activation shall be possible via hand switch and/or foot switch control to allow flexible surgeon operation during open and laparoscopic procedures. |  |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |  |
| - Each system shall include: - Ultrasonic generator unit – 1 - Compatible ultrasonic handpiece – minimum 2 - Footswitch – 1 - Handpiece cable – 1 - Starter set sufficient for initial clinical operation (if disposable type used) - Power cable – 1 - User and service manuals – 1 set |  |  |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |  |  |
| - Ultrasonic handpieces - Vessel sealing handpieces - Additional cables |  |  |  |  |  |  |  |
| - Operational Reliability [MUST] |  |  |  |  |  |  |  |
| - Stable operation during prolonged surgical procedures - Low acoustic and mechanical vibration outside operating tip - Rapid readiness after activation |  |  |  |  |  |  |  |
| - Installation & Acceptance Testing [MUST] |  |  |  |  |  |  |  |
| - Acceptance testing shall include: - Energy activation verification - Cutting and coagulation performance demonstration - Safety shutdown function verification - Alarm and status display verification - Acceptance checklist signed |  |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |  |
| - Clinical user training for surgeons and OR staff - Instrument handling and safety training - Basic biomedical maintenance orientation |  |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |  |
| - Minimum 2-year comprehensive warranty - Preventive maintenance included during warranty - Response time ≤ 72 hours - Spare parts availability ≥ 5–7 years |  |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-2 (HF surgical equipment safety or equivalent applicable energy device standard) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |

/End/