MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.27

| Priority | 1 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Thulium Fiber Laser | Installation Location | Operation Room 3 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The Thulium Fiber laser system shall be used for urological and endoscopic surgical procedures including lithotripsy, stone fragmentation, soft tissue incision, ablation, and coagulation. The system shall provide stable energy delivery with precise control and safe operation for minimally invasive surgical applications. - The system shall support integration with endoscopic surgical workflows, including compatibility with endoscopic visualization and irrigation systems used in urological procedures. - As the laser will be used in a multi-purpose operating room environment, the system shall incorporate appropriate laser safety provisions suitable for operating room use, including controlled activation and coordination with applicable operating room laser safety measures to ensure safe clinical operation. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Thulium laser generator - Mobile cart-based design suitable for OR use - Integrated cooling system (air or water cooling acceptable) - Compact footprint and easy maneuverability - Easy-to-clean and disinfectant-resistant surfaces - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Laser Output & Performance [MUST] |  |  |  |  |  |
| - Wavelength: approximately 1940 nm - Adjustable Pulse energy range approximately 0.02–6 J or wider. - Adjustable repetition rate: minimum 1–2400 Hz or higher - Maximum average power ≥ 60 Watts (or higher) - Stable and consistent pulse energy delivery - Pulse duration optimized for lithotripsy and tissue cutting - Multiple pulse modulation modes optimized for stone fragmentation and dusting techniques shall be available (manufacturer equivalent technology acceptable). - The system shall support both fragmentation and dusting surgical techniques for endourological procedures. |  |  |  |  |  |
| - Control & Interface [MUST] |  |  |  |  |  |
| - Digital control panel or touchscreen - Adjustable parameters: energy, frequency, power - Clear display of current settings - Ready/standby indication - Error and fault messages displayed |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Key switch activation - Emergency stop button - Audible and visual warning indicators during laser emission - Interlock for footswitch activation - Protective eyewear provided - Over-temperature protection - Fiber break detection or equivalent safety mechanism |  |  |  |  |  |
| - Delivery System & Accessories [MUST] |  |  |  |  |  |
| - Each system shall include: - Laser main unit – 1 - Footswitch – 1 - Laser fibers suitable for urology (minimum starter pack of 5 fibers with mixed sizes) - Fiber connectors/adapters – 1 set - Fiber stripping/cleaving tools – 1 set - Protective laser safety goggles – minimum 2 pairs - Mobile trolley/cart – 1 - Power cable – 1 - User and service manuals – 1 set - The system shall be compatible with standard reusable or disposable laser fibers from manufacturer-approved equivalent suppliers. All items required for immediate clinical operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional laser fibers and consumables - Fiber sterilization trays - Smoke evacuation system - Additional footswitch |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by certified personnel - Laser safety verification and calibration - Functional testing on-site - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for surgeons, nurses, and technicians - Training includes operation, safety procedures, fiber handling, and maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60825 (laser safety) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/