MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.28

| Priority | 1 | Floor | 1F-DR, 2F(Ward), 3F(Ward) | Quantity | 104EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Patient Bed | Installation Location | Labor Room: 4EA 2F(Ward): 50EA 3F(Ward): 50EA |  |
| Specification Level |  | Low | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The patient bed shall be used for general inpatient care, providing safe and comfortable positioning of adult patients. - The bed shall allow manual adjustment of backrest and leg sections and support routine nursing care. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Manual hospital bed with mechanical crank operation - Heavy-duty steel frame construction - Powder-coated or corrosion-resistant finish - Smooth surfaces for easy cleaning and disinfection - Stable and durable structure - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Dimensions & Load Capacity [MUST] |  |  |  |  |  |
| - Standard adult size bed (approx. 2000 mm × 900 mm mattress platform or equivalent) - Safe working load ≥ 200 kg - Mattress platform with perforation or ventilation slots |  |  |  |  |  |
| - Position Adjustment [MUST] |  |  |  |  |  |
| - Manual backrest adjustment (minimum 0–70°) - Manual knee/leg section adjustment - Crank handles foldable and smooth operation - Position locks secure under load |  |  |  |  |  |
| - Side Rails & Safety [MUST] |  |  |  |  |  |
| - Two or four side rails - Foldable or drop-down type - Secure locking mechanism - Rounded edges to prevent injury |  |  |  |  |  |
| - Mobility [MUST] |  |  |  |  |  |
| - Four caster wheels - Wheel diameter suitable for hospital flooring - Central or individual braking system - Easy maneuverability |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each bed shall include: - Bed frame – 1 - Side rails – 1 set - IV pole – 1 - Mattress (foam type, hospital grade) – 1 - Headboard and footboard – 1 set - User manual – 1 All items required for immediate patient use must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Over-bed table - Bedside locker - Pressure-relief mattress - Additional IV pole |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and assembly on-site - Inspection for safe operation |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - Comprehensive user training on operational functions, routine checks, and maintenance. - Maintenance training for technicians |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year warranty - Covers structural defects and mechanical parts - Spare parts availability ≥ 5 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - ISO or CE compliant hospital furniture standards - ISO 13485 or ISO 9001 manufacturer quality certification |  |  |  |  |  |

/End/