MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.3

| Priority | 1 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Bio Safety Cabinet | Installation Location | Clinical Pathology Laboratory |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The Biological Safety Cabinet (BSC) shall be used in the clinical pathology laboratory to provide personnel, product, and environmental protection during handling of infectious or potentially hazardous biological materials. - Equivalent or higher performance systems are acceptable. Specifications shall be vendor-neutral and performance-based. |  |  |  |  |  |
| - Classification & Compliance [MUST] |  |  |  |  |  |
| - Class II Type A2 Biological Safety Cabinet - Certified in accordance with EN 12469, NSF/ANSI 49, or equivalent international standard - Electrical safety compliant with IEC 61010 or IEC 60601 where applicable - Factory test certificate and airflow test report required - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Airflow & Filtration [MUST] |  |  |  |  |  |
| - Vertical laminar airflow design - HEPA filter efficiency ≥ 99.97% at 0.3 µm (supply and exhaust) - Average inflow velocity compliant with international biosafety standards (approx. 0.45 m/s) - Average downflow velocity compliant with international biosafety standards (approx. 0.30 m/s) - Negative pressure plenum or equivalent containment structure - Visual and audible alarms for airflow failure and unsafe sash position - Sash: Laminated multi-layer safety glass (UV resistant). - Interface & Functions - Display: Digital LCD showing air velocity, time, and status. - Monitoring: Airflow (Inflow/Downflow), UV Timer, Blower hour meter, Filter life indicator. - Security: Password protection for blower and service functions. - Connectivity: RS 232 or equivalent data output port. |  |  |  |  |  |
| - Construction & Materials [MUST] |  |  |  |  |  |
| - Work zone made of stainless steel (grade 304 or equivalent), seamless or easy-to-clean design - Corrosion-resistant external coating - Laminated safety glass front sash - Interior lighting (LED type) - Noise level suitable for laboratory environment (≤ 65 dBA) |  |  |  |  |  |
| - Dimensions [MUST] |  |  |  |  |  |
| - Nominal size: 4 ft (approximately 1.2 m width) - Adequate internal working height and depth for routine laboratory procedures |  |  |  |  |  |
| - Control & Monitoring [MUST] |  |  |  |  |  |
| - Microprocessor-based control system - Digital display for airflow status and filter condition - UV timer (with UV lamp supplied) - Blower hour meter or filter life indicator |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| All components required for safe and full operation must be included - Main cabinet unit - HEPA filters (installed) - Base stand with leveling feet (or fixed stand) - Internal LED lighting - UV lamp (interlocked with blower and sash) - Electrical outlet inside work zone (minimum 2) - Power cable - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| The following items shall NOT be included in the base price and must be quoted separately: - Gas/Vacuum/Air service valves - Additional electrical outlets - External exhaust connection kit - Annual re-certification service |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - On-site airflow validation and certification after installation - Integrity testing of HEPA filters - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for laboratory staff - Training on safe operation, cleaning, UV use, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance during warranty included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - Airflow performance certification - NSF/EN/IEC equivalent safety standards - CE marking or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/