MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 30

| Priority | 2 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Centrifuge (Refrigerated, Angle rotor) | Installation Location | Clinical Pathology Laboratory |  |
| Specification Level |  | Mid | Remarks |  |  |
| - General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The refrigerated centrifuge shall be used for separation of blood components, serum, plasma, and other biological samples requiring temperature-controlled centrifugation. - The system shall provide precise speed control, temperature stability, and safe operation for routine clinical laboratory use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Bench-top or floor-standing refrigerated centrifuge - Angle rotor configuration - Microprocessor-controlled operation - Quiet and low-vibration design - Easy-to-clean chemical-resistant housing - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Speed & Performance [MUST] |  |  |  |  |  |
| - Maximum speed ≥ 15,000 rpm - Maximum relative centrifugal force (RCF) ≥ 20,000 × g - Adjustable speed and RCF control - Fast acceleration and deceleration - Automatic imbalance detection and shutdown |  |  |  |  |  |
| - Temperature Control [MUST] |  |  |  |  |  |
| - Refrigeration system integrated - Temperature setting range: approximately -10°C to +40°C - Temperature stability during run - Pre-cooling function - Temperature display on control panel |  |  |  |  |  |
| - Rotor & Capacity [MUST] |  |  |  |  |  |
| - Angle rotor included as standard - Rotor capacity suitable for common laboratory tubes (e.g., 12–24 tubes, 5–15 mL size or equivalent) - Rotor material resistant to corrosion - Secure rotor locking mechanism |  |  |  |  |  |
| - Control & Safety Features [MUST] |  |  |  |  |  |
| - Digital display of speed, RCF, time, and temperature - Programmable run time - Lid lock during operation - Emergency lid release - Over-speed and over-temperature protection |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Centrifuge main unit – 1 - Angle rotor – 1 - Tube adapters as required for supplied rotor - Power cable – 1 - User manual and service manual – 1 set All components required for immediate laboratory operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional rotors (swing-out or different capacity) - Additional tube adapters - Aerosol-tight rotor lid |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and installation on laboratory bench or floor - Functional testing of speed and temperature - Safety verification prior to acceptance |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for laboratory staff - Training includes operation, cleaning, and routine maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 61010-1 (Safety requirements for laboratory electrical equipment) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/