MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 31

| Priority | 2 | Floor | GF, 2F, 3F | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Auto Height Weight Machine | Installation Location | Surgery Examination Room (1), Treatment Room (2) |  |
| Specification Level |  | Low | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The height and weight measuring system shall be used for routine assessment of patient body weight, height, and body mass index (BMI) in outpatient, emergency, and inpatient areas. - The system shall support high body weight (bariatric) patients and provide accurate, fast, and reliable measurements suitable for daily hospital use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Digital electronic weighing scale with integrated height measurement - Automatic BMI calculation function - Stable platform suitable for adult and bariatric patients - Heavy-duty construction for continuous hospital use - Easy-to-clean and disinfectant-resistant surface - Anti-slip platform surface - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Weight Measurement Performance [MUST] |  |  |  |  |  |
| - Maximum capacity ≥ 300 kg - Graduation/accuracy ≤ 0.1 kg (or better) - High precision load cells - Fast stabilization time - Zero/tare function - Overload protection - Automatic hold function (to capture unstable patient readings) |  |  |  |  |  |
| - Height Measurement Performance [MUST] |  |  |  |  |  |
| - Integrated stadiometer/height rod - Height measurement range suitable for adult patients (minimum approx. 60–210 cm or wider) - Graduation ≤ 0.5 cm - Easy manual or automatic adjustment |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |
| - Digital display clearly visible to operator - Simultaneous display of weight and height - Automatic BMI calculation and display - Simple control buttons - Backlit display for low-light areas |  |  |  |  |  |
| - Platform & Mechanical Design [MUST] |  |  |  |  |  |
| - Wide and low-profile platform for easy patient access - Platform size sufficient for large patients - Non-slip surface - Stable anti-tip structure - Optional handrail compatibility - Wheelchair-accessible design acceptable if offered |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC power 100–240 V, 50/60 Hz - Rechargeable battery backup for portable use - Battery operation ≥ 8 hours minimum - Automatic charging |  |  |  |  |  |
| - Data & Connectivity [MUST] |  |  |  |  |  |
| - Internal memory or last value recall - USB or equivalent data output capability (if available) - Date/time stamp optional but acceptable |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Digital scale main unit – 1 - Integrated height rod – 1 - Power adapter/cable – 1 - Rechargeable battery (installed) - User manual – 1 All components required for immediate clinical operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Handrail support - Printer - BMI report printing module - Wheelchair ramp - Data management software |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and on-site setup - Calibration verification - Functional testing prior to handover |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - Basic user training for nursing and OPD staff - Instruction on calibration and cleaning procedures |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year comprehensive warranty - Covers parts, labor, and service visits - Response time ≤ 48 hours - Spare parts availability ≥ 5 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - CE marking or equivalent international regulatory approval for medical or measuring devices - Accuracy shall comply with applicable international weighing instrument standards (e.g., OIML R76 or equivalent) - Manufacturer shall be ISO 9001 or ISO 13485 certified. |  |  |  |  |  |

/End/