MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 32

| Priority | 2 | Floor | GF, 1F-DR, 2F, 3F | Quantity | 4EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Patient Monitor (Portable) | Installation Location | OB-GYN Treatment Room (1), Labor Room (1), Treatment Room (2) |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The portable patient monitor shall be used for continuous monitoring of vital signs during bedside care and intra-hospital transport. - The system shall be suitable for adult, pediatric, and neonatal patients in emergency, labor, and treatment environments. - The monitor shall support rapid deployment, reliable transport operation, and safe alarm management. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Compact portable monitor suitable for bedside and transport use, with integrated carry handle and bed hook - Rugged construction and shock-resistant design appropriate for transport within hospital - Easy-to-clean, disinfectant-resistant housing and smooth surfaces - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. - On-screen language: English - Capable of continuous operation (24/7 clinical use) |  |  |  |  |  |
| - Parameters Included in Base Supply [MUST] |  |  |  |  |  |
| - ECG with Heart Rate - Respiration rate (derived from ECG/impedance or equivalent) - SpO₂ with Pulse Rate - Non-invasive blood pressure (NIBP) with SYS/DIA/MAP - Temperature (minimum 1 channel) |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Color display size ≥ 8 inches - Simultaneous display of at least 3 waveforms and numeric values - Adjustable brightness for day/night use - Touchscreen or knob/keypad control (either acceptable), with fast parameter access - On-screen battery status, alarm status, and patient type selection (adult/ped/neonate) |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |
| - ECG - 3-lead and/or 5-lead monitoring (either acceptable; bidder to specify) - Heart rate range: at least 20–300 bpm - Arrhythmia detection and alarm capability - Pacemaker detection (if available) acceptable - Respiration - Respiration range: at least 0–120 breaths/min - Apnea alarm capability - SpO₂ - SpO₂ range: 0–100% - Accuracy: ±2% (70–100%) or better (manufacturer standard) - Motion-tolerant performance (technology unspecified; performance-based) - NIBP - Oscillometric method - Patient modes: Adult / Pediatric / Neonatal - Measurement range suitable for adult/pediatric/neonate clinical use - Accuracy compliant with ISO 81060-2 / AAMI or equivalent standard - Auto-interval measurement: adjustable 1–60 minutes (or wider) - STAT mode (rapid repeated readings) or equivalent function - Temperature - Compatible with standard skin or probe temperature sensor - Measurement range suitable for clinical use (approx. 25–45°C) |  |  |  |  |  |
| - Alarms & Safety [MUST] |  |  |  |  |  |
| - Audible and visual alarms for all parameters - Adjustable alarm limits and alarm volume - Alarm priority differentiation (high/medium/low) or equivalent - Sensor disconnection alarms (ECG/SpO₂/NIBP/Temp) - Battery low and power failure alarms - Alarm silence with automatic re-alarm (time-limited) or equivalent safety feature |  |  |  |  |  |
| - Trends, Data & Connectivity [MUST] |  |  |  |  |  |
| - Trend storage: minimum 24 hours of numeric trends (or better) - Event log including alarms and parameter limit violations - Data export via USB and/or LAN (either acceptable; bidder to specify) - Capability for connection to central monitoring system (if available). - Date/time stamp capability |  |  |  |  |  |
| - Power & Portability [MUST] |  |  |  |  |  |
| - AC power: 100–240V, 50/60Hz - Internal rechargeable battery backup: ≥ 4 hours continuous monitoring - Quick battery swap or serviceable battery design acceptable - Weight and form factor suitable for transport (portable bedside use) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Monitor main unit – 1 - ECG patient cable with lead wires – 1 set - SpO₂ reusable probe (adult/pediatric) – 1 - SpO₂ neonatal probe – 1 - NIBP hose – 1 - NIBP cuffs: adult (1), pediatric (1), neonatal (1) – 1 set each - Temperature probe/sensor – 1 - Battery (installed) – 1 - Power adapter/cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - EtCO₂ module (mainstream) with sampling lines - Additional probes/cuffs/consumables - Portable trolley/stand - Carry bag or protective case |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and on-site commissioning by qualified personnel - Functional testing of parameters and alarms - Electrical safety test verification - Acceptance testing checklist provided and completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for nurses and clinicians - Training includes alarm management, cleaning, and routine checks - Basic troubleshooting guidance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance at least once per year during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (medical electrical safety) and IEC 60601-1-2 (EMC) or equivalent - IEC 60601-2-27 (ECG monitoring), IEC 60601-2-49 (multiparameter monitors) or equivalent (as applicable) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/