MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 35

| Priority | 2 | Floor | 1F-DR | Quantity | 4EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Patient Monitor (Console-NICU) | Installation Location | NICU |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The bedside patient monitor shall be used for continuous monitoring of patients in ward and NICU environments requiring extended monitoring duration and centralized observation capability. - The system shall support adult, pediatric, and neonatal monitoring with reliable alarm management and networking capability. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Bedside multiparameter monitor designed for continuous 24/7 use - Modular or configurable architecture to support parameter expansion - Durable, disinfectant-resistant housing suitable for clinical environment - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. - On-screen language: English |  |  |  |  |  |
| - Parameters Included in Base Supply [MUST] |  |  |  |  |  |
| - ECG with Heart Rate - Respiration rate - SpO₂ with Pulse Rate - NIBP with SYS/DIA/MAP - Temperature (minimum 1 channel) - IBP x 2 |  |  |  |  |  |
| - Expandability Requirements [MUST] |  |  |  |  |  |
| The monitor shall be capable of adding the following modules (built-in or plug-in; bidder to specify): - Invasive blood pressure (IBP): at least 2 channels capability - EtCO₂ (capnography) capability - (Modules may be supplied as Optional if not included in base; see Optional section.) |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Color display size ≥ 12 inches - Display of at least 6 waveforms simultaneously - Numeric and graphical trend display - Touchscreen and knob control - Adjustable brightness for day/night - Patient type selection: adult / pediatric / neonatal |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |
| - ECG - 3/5 lead monitoring (bidder to specify) - Heart rate range: at least 20–300 bpm - Arrhythmia detection and alarm capability - Respiration - Respiration range: at least 0–120 breaths/min - Apnea alarm capability - SpO₂ - SpO₂ range: 0–100% - Accuracy: ±2% (70–100%) or better - Motion-tolerant performance - NIBP - Oscillometric method - Adult / Pediatric / Neonatal modes - Accuracy compliant with ISO 81060-2 / AAMI or equivalent - Auto-interval: adjustable 1–60 minutes (or wider) - STAT mode or equivalent - Temperature - Measurement range suitable for clinical use (approx. 25–45°C) - IBP - Measurement range: 0-300mmHg - Parameters: Systolic, Diastolic, Mean Arterial Pressure, Pulse Rate - Adult/Pediatric/Neonatal modes - Trends - Trend storage: minimum 72 hours numeric trends - Event log and alarm log storage The monitor shall be suitable for neonatal monitoring applications including low-perfusion and motion-tolerant SpO₂ measurement performance. |  |  |  |  |  |
| - Alarms & Safety [MUST] |  |  |  |  |  |
| - Audible and visual alarms for all parameters - Adjustable alarm limits and volume - Alarm priority differentiation (high/medium/low) or equivalent - Sensor disconnection alarms - Battery low and power failure alarms - Alarm silence with automatic re-alarm (time-limited) or equivalent |  |  |  |  |  |
| - Networking & Data [MUST] |  |  |  |  |  |
| - LAN connectivity (wired) for central monitoring system connection - Standard protocols or manufacturer solution acceptable (no brand restriction) - USB data export - Central monitoring station compatibility (bidder to specify solution) - Date/time stamp and patient ID entry capability |  |  |  |  |  |
| - Power [MUST] |  |  |  |  |  |
| - AC power: 100–240V, 50/60Hz - Internal rechargeable battery backup: ≥ 2 hours |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical operation: - Monitor main unit – 1 - ECG patient cable with lead wires – 1 set - SpO₂ reusable probe (adult/pediatric) – 1 - SpO₂ neonatal probe – 1 - NIBP hose – 1 - NIBP cuffs: adult (1), pediatric (1), neonatal (1) – 1 set each - Temperature probe/sensor – 1 - IBP transducers and pressure tubing sets (starter packs) - Battery (installed) – 1 - Power cable – 1 - Mounting: wall mount or trolley/stand – 1 (as appropriate for ward/NICU use) - User manual and service manual – 1 set |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - EtCO₂ module (side stream or mainstream) and sampling lines - Central monitoring station and display - Additional probes/cuffs/consumables - Wi‑Fi module (if offered) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and installation/commissioning by qualified personnel - Configuration for ward/NICU use (alarm defaults per patient type) - Functional testing of parameters and alarms - Electrical safety test verification - Acceptance testing checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for nurses and clinicians - Training includes alarm management, cleaning, daily checks, and connectivity overview - Basic troubleshooting and escalation procedure |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs, and included accessories - Response time ≤ 48 hours - Preventive maintenance at least once per year during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (medical electrical safety) and IEC 60601-1-2 (EMC) or equivalent - IEC 60601-2-27 (ECG monitoring), IEC 60601-2-49 (multiparameter monitors) or equivalent (as applicable) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified |  |  |  |  |  |

/End/