MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 36

| Priority | 2 | Floor | GF, 1F-DR | Quantity | 2EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Electric Examination Table | Installation Location | Obstetrics and Gynecology Examination Room (1) Labor Room (1) |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The electric examination table shall be used for OB-GYN examinations and minor procedures, providing safe patient positioning, clinician ergonomics, and infection-control-friendly surfaces. - The table shall support smooth electric positioning for height and backrest/seat adjustments and include OB-GYN accessories. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electrically powered (electric or electro-hydraulic) examination table - Smooth, quiet motion with stable locking under load - Robust frame suitable for continuous outpatient clinical use - Easy-to-clean, disinfectant-resistant upholstery and surfaces - Rounded edges and safe design to prevent injury - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. - On-product labels/manuals in English |  |  |  |  |  |
| - Dimensions & Load Capacity [MUST] |  |  |  |  |  |
| - Standard adult examination table size suitable for OB-GYN exams - Safe working load ≥ 200 kg - Stable base with anti-tip design |  |  |  |  |  |
| - Positioning Functions [MUST] |  |  |  |  |  |
| - Electric height adjustment - Electric backrest adjustment - Seat/leg section positioning suitable for gynecological examination posture - Trendelenburg function (manual or electric) preferred, minimum ±10° (if offered) - Smooth start/stop control to avoid patient discomfort |  |  |  |  |  |
| - Height & Motion Range [MUST] |  |  |  |  |  |
| - Height range: approximately 480–1000 mm or wider - Backrest adjustment: at least 0–60° - Foot section/leg section adjustment suitable for lithotomy positioning (mechanical or assisted acceptable) |  |  |  |  |  |
| - Control System & Safety [MUST] |  |  |  |  |  |
| - Hand control and/or footswitch control included - Low-voltage control system with safety protections - Emergency stop function or equivalent safety cutoff - Anti-crush protection or safe stop function (if available) |  |  |  |  |  |
| - Patient Comfort & Infection Control [MUST] |  |  |  |  |  |
| - High-density foam mattress with seamless or minimal-seam upholstery - Upholstery resistant to common hospital disinfectants - Removable/cleanable cushions where applicable |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - The supply shall include all items required for immediate clinical use: - Electric examination table main unit – 1 - Hand control and/or footswitch – 1 - Leg holders/stirrups – 1 set - Arm supports (if standard) – 1 set - Stainless steel bowl or drainage tray – 1 - Paper roll holder or equivalent (if offered) – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Examination lamp attachment - Additional leg holder sets or padding - Protective covers and accessories - Stool/chair (if offered) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and on-site setup - Functional testing of all movements and safety features - Electrical safety verification - Acceptance checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for clinicians and nurses - Cleaning and routine checks instruction |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, travel costs - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/