MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.37

| Priority | 2 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Electro Surgical Unit | Installation Location | ER Treatment Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The electrosurgical unit shall be used in the emergency department for cutting and coagulation during minor surgical procedures, wound management, bedside interventions, and emergency hemostasis procedures. - The ESU shall provide safe and stable monopolar and bipolar performance with appropriate safety monitoring to reduce risk of patient burns and procedural complications. - Equivalent or higher performance systems are acceptable. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled electrosurgical generator for clinical use - Monopolar cutting and coagulation functions - Bipolar coagulation function - Suitable for frequent ER use with fast setup - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Output Performance – Monopolar [MUST] |  |  |  |  |  |
| - Cutting modes: pure cut and blended cut (minimum) - Coagulation modes: contact coagulation and spray coagulation (minimum) - Output power (monopolar cut): ≥ 300 W - Output power (monopolar coag): ≥ 200 W - Output power adjustment with clear digital indication - Stable output under variable tissue impedance |  |  |  |  |  |
| - Output Performance – Bipolar [MUST] |  |  |  |  |  |
| - Bipolar coagulation mode suitable for ER procedures - Output power (bipolar): ≥ 80–100 W - Fine power adjustment and stable performance |  |  |  |  |  |
| - Safety Monitoring [MUST] |  |  |  |  |  |
| - Return electrode monitoring (REM) or equivalent patient plate contact quality monitoring - Audio/visual alarms for return electrode fault - Overload/overtemperature protection - Self-test at startup - Error code display |  |  |  |  |  |
| - User Interface & Controls [MUST] |  |  |  |  |  |
| - Digital display of selected mode and power - Easy selection of modes suitable for ER staff workflow - Footswitch and/or hand-switch compatibility - Audible activation tone |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate use: - ESU main generator – 1 - Monopolar handpiece with electrodes (starter set) – 1 set - Reusable patient return electrode or disposable type with compatible cable – 1 set - Bipolar forceps (reusable) – 1 - Bipolar cable – 1 - Monopolar Footswitch (dual pedal) – 1 - Bipolar footswitch-1 - Power cable – 1 - User manual and service manual – 1 set The electrosurgical unit shall provide a smoke evacuation interface or compatibility with external smoke evacuation systems. |  |  |  |  |  |
| - Electrical Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Proper grounding requirement specified |  |  |  |  |  |
| - Infection Control & Durability [MUST] |  |  |  |  |  |
| - Sealed/wipeable front panel - Disinfectant-resistant casing - Durable connectors designed for frequent plug/unplug |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - On-site installation and functional verification - Verification of monopolar and bipolar modes - Alarm/REM function test - Acceptance checklist signed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - User training for ER doctors and nurses - Includes safe use, plate placement, mode selection, and troubleshooting - Biomedical orientation for routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Warranty minimum 3 years after acceptance - PM included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-2 (Particular requirements for HF surgical equipment) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/