MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.38

| Priority | 2 | 3 | Floor | GF, 1F-DR | Quantity | 2EA | 1EA |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  | Operating Light (Portable) | Installation Location | ER Treatment Room (1) |  |  |
|  |  |  |  |  | Delivery Room (1) |  |  |
|  |  |  |  |  | Obstetrics and Gynecology Clinic Room (1) |  |  |
| Specification Level |  |  | Mid | Remarks |  |  |  |
| General Specification |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |
| - The portable operating light shall provide high-intensity, shadow-reduced illumination suitable for minor surgical procedures, emergency interventions, and obstetric or examination procedures in emergency, treatment, and delivery room environments. - The light shall ensure adequate brightness, uniform field illumination, and flexible positioning with stable locking to support safe clinical procedures. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |
| - Mobile surgical/procedure light on stable wheeled stand - LED light source preferred for low heat and long life - Adjustable intensity and positioning - Disinfectant-resistant surfaces - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |
| - Illumination Performance [MUST] |  |  |  |  |  |  |  |
| - Light intensity at 1 m distance: ≥ 120,000 lux (or higher) - Adjustable intensity with multiple steps or continuous dimming - Light field diameter: approximately 18–25 cm (adjustable preferred) - Color temperature: approximately 4,000–5,000 K (neutral white) - Color rendering index (CRI): ≥ 95 - Low shadow effect with good depth illumination |  |  |  |  |  |  |  |
| - Mechanical & Positioning [MUST] |  |  |  |  |  |  |  |
| - Multi-joint arm allowing flexible positioning - Smooth movement with secure lock/brake - Stable stand with anti-tip design - Lockable caster wheels - Easy maneuverability in ER corridors/rooms |  |  |  |  |  |  |  |
| - Thermal & Acoustic Performance [MUST] |  |  |  |  |  |  |  |
| - Low heat emission at surgeon head area - Noise level suitable for clinical environment (low-noise operation) |  |  |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Battery backup capability is acceptable to support temporary operation during patient movement or power interruption. |  |  |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |  |  |
| - Electrical safety compliant design - Overheat protection (for LED driver) - Stable structure to prevent accidental fall - Cable management to prevent tripping |  |  |  |  |  |  |  |
| - Accessories [MUST] |  |  |  |  |  |  |  |
| - Light head with handle (sterilizable or disposable covers) - Mobile stand with arms – 1 - Sterile handle covers – minimum 20 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |  |  |
| - Functional test of intensity and dimming - Mechanical movement and brake test - Acceptance checklist signed |  |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |  |
| - User training for staff including positioning and cleaning |  |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |  |
| - Warranty minimum 2 years - Spare parts availability ≥ 5–7 years |  |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |

/End/