MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 39

| Priority | 2 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Fetal Doppler | Installation Location | Obstetrics and Gynecology Clinic Room |  |
| Specification Level |  | Low | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The handheld fetal Doppler shall be used for non-invasive detection and intermittent monitoring of fetal heart rate during antenatal examination, labor assessment, and bedside obstetric evaluation. - The device shall provide accurate fetal heart rate detection with clear audio output and numeric display to support routine clinical assessment. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Portable handheld Doppler unit - Lightweight and ergonomic design - Durable housing suitable for daily clinical use - Easy-to-clean surface - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |
| - Ultrasound frequency suitable for obstetric fetal heart rate detection (typically around 2–3 MHz). - Fetal heart rate measurement range approx. 50–210 bpm - Digital heart rate display - Clear audible heart sound output - Fast detection capability |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |
| - Numeric LCD/LED display of fetal heart rate - Volume control - Power/battery indicator |  |  |  |  |  |
| - Probe Requirements [MUST] |  |  |  |  |  |
| - Waterproof or gel-resistant probe - Replaceable probe or cable if damaged - Durable cable strain relief - Probe cable shall include reinforced strain relief design suitable for frequent clinical handling. |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - Battery operated (rechargeable) - Minimum 4–6 hours continuous operation - Low battery warning - Auto power-off function to preserve battery life is acceptable. |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Doppler main unit – 1 - Probe – 1 - Rechargeable battery – 1 set - Charger – 1 - Carry pouch – 1 - User manual – 1 |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional probes |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |
| - Delivery and functional testing - Basic user training for nurses/midwives |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year warranty - Spare parts availability ≥ 5 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-37 (Particular requirements for ultrasonic medical diagnostic and monitoring equipment) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/