MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.4

| Priority | 1 | Floor | GF | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Defibrillator | Installation Location | Labor Room (1), Treatment Room (2) |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The defibrillator/monitor shall be used in emergency, treatment, and labor rooms for cardiac resuscitation, cardioversion, pacing, and patient monitoring. - The system shall be portable, reliable, and suitable for rapid emergency response. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Portable biphasic defibrillator/monitor - Manual defibrillation + AED mode - Synchronized cardioversion - Non-invasive pacing - Integrated patient monitoring (ECG and SpO₂ minimum) - Durable construction suitable for emergency use - Fast start-up and charging - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Clinical Functions [MUST] |  |  |  |  |  |
| - Defibrillation - Biphasic waveform technology - Energy selection adjustable up to ≥ 200 Joules - Fast charge time ≤ 5 seconds to typical shock energy - Manual and AED modes - Visual and voice prompts in AED mode - Cardioversion - Synchronized cardioversion capability - Pacing - Non-invasive transcutaneous pacing - Adjustable pacing rate and output - Demand and fixed modes - Monitoring - ECG monitoring (pads/paddles and 3 or 5-lead cable) - SpO₂ monitoring integrated - Audible and visual alarms |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Color display ≥ 6 inches - Clear waveform and numeric display - Simple and intuitive controls suitable for emergency use - Status indicators for readiness and battery level |  |  |  |  |  |
| - Power & Portability [MUST] |  |  |  |  |  |
| - Rechargeable battery powered - Continuous monitoring/operation time ≥ 2 hours minimum - AC operation with automatic charging - Universal power supply (100–240V, 50/60Hz) - Lightweight portable design - Splash resistance protection (IPX4 or higher) |  |  |  |  |  |
| - Recording & Data [MUST] |  |  |  |  |  |
| - Internal memory for event recording - Integrated thermal printer - USB or equivalent data export |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Main defibrillator/monitor unit - Rechargeable battery × 1 (installed) - Spare battery × 1 - Adult/pediatric multifunction pads or paddles – 1 set - ECG cable – 1 set - SpO₂ sensor (reusable) – 1 set - Printer paper – starter pack - Test load – 1 - Carry case or protective cover - Power cable - All required software and licenses - User and service manuals |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional batteries - Additional sensors or disposable pads - External mounting brackets or carts |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional and safety testing on-site - Device supplied fully ready for use - Acceptance testing included |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for clinical staff - Training on defibrillation, pacing, monitoring, and routine checks |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, battery, paddles, and accessories - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-2-4 - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/