MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.40

| Priority | 2 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Cast Removal System | Installation Location | Orthopedic Treatment Room |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The cast removal system shall be used for safe and efficient removal of plaster and synthetic casts in orthopedic, emergency, and outpatient clinical settings. - The system shall minimize risk of patient skin injury through controlled oscillating cutting action and shall include dust collection capability to reduce airborne particulate exposure during cast removal procedures. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade oscillating cast cutter system - Suitable for plaster of Paris and fiberglass/synthetic casts - Ergonomic handpiece for repeated daily use - Integrated dust extraction system OR compatibility with external vacuum system shall be provided. - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Cutting Performance [MUST] |  |  |  |  |  |
| - Oscillating cutting mechanism (NOT rotary) - Adequate power for thick casts without stalling - Variable speed control or optimized speed suitable for cast removal - Low vibration design to improve control |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Blade guard and safe blade mounting mechanism - Overheat protection - Low noise design suitable for clinic use - Reduced skin heat generation during prolonged cutting - Quick stop function - The system shall minimize blade heat transfer to the patient surface during prolonged operation to reduce risk of skin injury. |  |  |  |  |  |
| - Dust Collection [MUST] |  |  |  |  |  |
| - The system shall include an integrated dust collection unit or verified compatibility with a dedicated external vacuum system suitable for cast removal applications. - Replaceable filter/bag system - Easy cleaning and maintenance |  |  |  |  |  |
| - Ergonomics & Mechanical [MUST] |  |  |  |  |  |
| - Lightweight handpiece with non-slip grip suitable for prolonged clinical use. - Blade replacement mechanism shall allow simple and safe replacement without requiring special tools or complex adjustment procedures. - Durable cable construction with adequate strain relief to withstand routine clinical handling. |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Electrical safety protection |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Cast cutter main unit/handpiece – 1 - Power supply/control unit (if separate) – 1 - Vacuum unit or dust collection module – 1 - Cast cutting blades (assorted) – minimum 10 - Blade wrench/tool (if required) – 1 - Filters/dust bags – minimum 10 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - Functional test including cast cutting demonstration on sample material - Vacuum/dust collection function verification - Acceptance checklist signed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for orthopedic staff and nurses - Safe technique, blade change, cleaning, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Warranty minimum 2 years - Response time ≤ 72 hours - Spare parts availability ≥ 5–7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/