MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 41

| Priority | 2 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Histology Tissue Processing System | Installation Location | Histopathology Lab |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The histology tissue processing system shall be used for automated processing of tissue specimens including dehydration, clearing, and paraffin infiltration prior to embedding for histopathological examination. - The system shall provide reliable, safe, and reproducible processing suitable for routine clinical laboratory operation. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Fully automated tissue processor - Closed or semi-closed reagent handling system - Microprocessor-controlled operation - Bench-top or floor-standing configuration - Chemical-resistant housing - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. - The system shall minimize operator exposure to reagent fumes through sealed or controlled reagent handling design. |  |  |  |  |  |
| - Processing Capacity [MUST] |  |  |  |  |  |
| - Processing capacity: minimum 80 cassettes per run or higher. - Multiple reagent stations (minimum 10–12 or more) - Paraffin wax stations included |  |  |  |  |  |
| - Processing Functions [MUST] |  |  |  |  |  |
| - Automated dehydration, clearing, and paraffin infiltration - Programmable processing cycles - Adjustable temperature and time settings - Overnight/extended run capability - Continuous or batch processing |  |  |  |  |  |
| - Temperature Control [MUST] |  |  |  |  |  |
| - Paraffin chamber temperature control - Stable and uniform heating - Over-temperature protection |  |  |  |  |  |
| - Reagent Management & Safety [MUST] |  |  |  |  |  |
| - Sealed or covered reagent containers to reduce fumes - Fume control or exhaust port - Spill-resistant design - Alarm for low reagent level or system error - Safety lock or cover during operation |  |  |  |  |  |
| - Control & Interface [MUST] |  |  |  |  |  |
| - Digital display with cycle status - Program memory for multiple protocols - Audible/visual alarms - User-friendly interface |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Main tissue processor – 1 - Reagent containers – full set - Paraffin containers – full set - Cassette baskets – minimum 2 - Forceps/tools – 1 set - Power cable – 1 - User and service manuals – 1 set All items necessary for immediate laboratory operation must be included. |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional baskets - Reagent containers - Fume extraction system |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and installation by qualified personnel - Functional testing of all stations - Temperature calibration |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for laboratory staff - Training includes programming, cleaning, and maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Preventive maintenance included - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 61010-1 (Safety requirements for laboratory electrical equipment) - CE or equivalent laboratory equipment approval - Manufacturer shall be ISO 13485 or ISO 9001 certified. |  |  |  |  |  |

/End/