MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 42

| Priority | 2 | Floor | 1F-DR, 2F(Ward), 3F(Ward) |  | Quantity | 12EA |
| --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Infusion Pump |  | Installation Location | Treatment Room (8), NICU (4) |  |
| Specification Level |  | Mid |  | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |
| - The infusion pump shall provide accurate and controlled intravenous infusion of fluids, medications, and parenteral solutions for adult, pediatric, and neonatal patients. - The device shall ensure safe delivery with precise flow control, stable low-flow performance suitable for neonatal applications, and comprehensive alarm functions. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |
| - Volumetric infusion pump - Compact portable design - Suitable for bedside and transport use - Durable and disinfectant-resistant housing - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |
| - Infusion Performance [MUST] |  |  |  |  |  |  |
| - Flow rate range: minimum 1–1200 mL/h or wider - Accuracy: ±5% or better - Programmable infusion volume and time - Bolus function - KVO (keep vein open) rate function - Suitable for adult, pediatric, and neonatal use |  |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |  |
| - Clear digital display - Flow rate, volume infused, and remaining time display - Simple keypad or touchscreen interface - Backlit display |  |  |  |  |  |  |
| - Safety & Alarm Functions [MUST] |  |  |  |  |  |  |
| - Occlusion alarm (adjustable sensitivity) - Air-in-line alarm - Door open alarm - Empty container alarm - Low battery alarm - End of infusion alarm - Audible and visual alarm indicators - The system shall minimize unintended bolus delivery following occlusion release. |  |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |  |
| - AC 100–240 V - Internal rechargeable battery - Battery backup ≥ 4 hours continuous operation |  |  |  |  |  |  |
| - IV Set Compatibility [MUST] |  |  |  |  |  |  |
| - Anti-free flow mechanism - Set loading safety design - The infusion pump shall be compatible with commonly available standard disposable IV sets without mandatory proprietary consumables. |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |
| - Each unit shall include: - Infusion pump main unit – 1 - IV pole clamp – 1 - Power cable – 1 - Starter IV sets – minimum 10 - User manual – 1 All components required for immediate clinical use must be included. |  |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |  |
| - Docking station - Additional IV sets - Drug library software (if available) |  |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |  |
| - Delivery and functional testing - Calibration verification |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |
| - On-site training for nursing staff - Programming and alarm management training |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-24 (Particular requirements for infusion pumps and controllers) - CE marking and FDA clearance or equivalent international regulatory approval |  |  |  |  |  |  |

/End/