MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 43

| Priority | 2 | Floor | 1F-DR, 1F-OT, 2F(Ward), 3F(Ward) |  | Quantity | 11EA |
| --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Syringe Pump |  | Installation Location | Treatment Room(2), NICU(4), Recovery Room(2), Operation Room(3) |  |
| Specification Level |  | Mid |  | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |
| - The syringe pump shall provide accurate and continuous administration of intravenous medications and parenteral solutions using disposable syringes for adult, pediatric, and neonatal patients. - The system shall ensure precise low-flow infusion suitable for neonatal and critical care applications with comprehensive safety alarm functions. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |
| - Microprocessor-controlled syringe infusion pump - Compact and lightweight design suitable for bedside and transport use - Durable, disinfectant-resistant housing - Quiet operation suitable for NICU environment - Easy-to-use interface for nursing staff - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |
| - Infusion Performance [MUST] |  |  |  |  |  |  |
| - Flow rate range: minimum 0.1–1200 mL/h or wider - Low-flow capability suitable for neonatal infusion (≤ 0.1 mL/h resolution) - Accuracy: Mechanical accuracy ≤ ± 1%; Overall accuracy (with syringe) ≤ ± 2%. - Micro mode: 0.1 – 99.9 mL/h with 0.01 mL/h increments. - Programmable infusion rate, volume, and time - Bolus function with adjustable bolus rate from 50 mL/h up to 1200 mL/h. - KVO (keep vein open) rate function. - Continuous and stable infusion without pulsation |  |  |  |  |  |  |
| - Syringe Compatibility [MUST] |  |  |  |  |  |  |
| - The system shall be compatible with commonly available standard disposable syringes without mandatory proprietary consumables. - Supported syringe sizes: minimum 5, 10, 20, 30, 50/60 mL - Automatic syringe size detection or easy manual selection - Anti-free-flow protection mechanism - Secure syringe locking system |  |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |  |
| - Digital backlit display - Display of: - Infusion rate - Volume infused - Volume remaining - Time remaining - Alarm status - Simple keypad or touchscreen interface - Clear parameter setting and confirmation |  |  |  |  |  |  |
| - Alarm & Safety Features [MUST] |  |  |  |  |  |  |
| - Audible and visual alarms for: - Occlusion - Air-in-line - End of infusion - Syringe empty - Door open - Low battery - System fault - Adjustable occlusion sensitivity - Automatic alarm logging - Self-test on startup - Anti-bolus protection after occlusion release - The system shall minimize unintended bolus delivery following occlusion release or system restart. |  |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |  |
| - AC power: 100–240 V, 50/60 Hz - Internal rechargeable battery - Battery backup ≥ 4–6 hours continuous operation - Battery level indicator |  |  |  |  |  |  |
| - Mounting & Mechanical Features [MUST] |  |  |  |  |  |  |
| - IV pole clamp included - Stackable or docking capability for multiple pumps (if available) - Stable mounting without drift - Lightweight and portable design |  |  |  |  |  |  |
| - Data & Drug Library [MUST] |  |  |  |  |  |  |
| - Event history or log storage - Dose/volume history review capability - Drug library function acceptable (not mandatory but allowed) - USB or equivalent service/data port |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Syringe pump main unit – 1 - IV pole clamp – 1 - Power adapter/cable – 1 - Rechargeable battery (installed) - Starter syringes or compatible sets (minimum 10) - User manual and service manual – 1 set All components required for immediate operation must be included. |  |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |  |
| - Docking station or multi-pump rack - Additional clamps - Additional syringe sets - Drug library software expansion |  |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |  |
| - Delivery and on-site setup - Functional testing and calibration verification - Electrical safety testing - Acceptance testing checklist completed |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |
| - On-site training for nurses and clinical staff - Programming and alarm management - Routine maintenance and cleaning instructions |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-24 (Particular requirements for infusion pumps and controllers) - CE marking and FDA clearance or equivalent international regulatory approval |  |  |  |  |  |  |

/End/