MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 44

| Priority | 2 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Ultrasound (OT Recovery Room) | Installation Location | Recovery Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The ultrasound system shall provide real-time diagnostic ultrasound imaging to support perioperative and postoperative patient assessment in recovery and critical care environments. - The system shall enable rapid bedside evaluation of cardiac function, lung condition, abdominal status, vascular access, and fluid responsiveness. - The system shall be suitable for point-of-care ultrasound (POCUS), anesthesia assessment, postoperative complication detection, and urgent clinical decision-making. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Cart-based high-performance diagnostic ultrasound system - Designed for continuous clinical use in perioperative environments - Rapid boot time suitable for emergency assessment - Medical-grade disinfectant-resistant surfaces - Silent cooling suitable for recovery room environment - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Imaging Modes [MUST] |  |  |  |  |  |
| - System shall support: - B-mode (2D imaging) - M-mode - Color Doppler - Power Doppler - Pulsed Wave Doppler (PW) - Tissue Doppler Imaging (TDI) - Harmonic imaging - Compound imaging - Speckle reduction imaging - Continuous Wave Doppler (CW) capability acceptable if available. - System shall allow simultaneous real-time imaging and Doppler analysis. |  |  |  |  |  |
| - Clinical Applications [MUST] |  |  |  |  |  |
| - System shall support multi-application perioperative use: - Cardiac: - Basic echocardiographic assessment - LV function evaluation - Pericardial effusion detection - Lung: - Pleural sliding assessment - Pneumothorax screening - B-line evaluation - Abdominal: - FAST examination capability - Fluid detection - Vascular: - Ultrasound-guided central and peripheral line placement - Venous thrombosis assessment - MSK / Soft Tissue: - Procedural guidance capability |  |  |  |  |  |
| - Transducers [MUST] |  |  |  |  |  |
| - System shall include minimum: - Phased array probe (cardiac application) - Convex probe (abdominal/general imaging) - Linear probe (vascular/procedural guidance) - Requirements: - Broadband multi-frequency probes - Electronic probe recognition - Probe temperature monitoring |  |  |  |  |  |
| - Image Quality Performance [MUST] |  |  |  |  |  |
| - High-resolution digital beamforming - Wide dynamic range suitable for diagnostic ultrasound imaging (manufacturer standard). - Real-time image optimization - Automatic image optimization function - Depth penetration suitable for adult patients - Frame rate optimized for cardiac imaging |  |  |  |  |  |
| - Measurement & Analysis Tools [MUST] |  |  |  |  |  |
| - Cardiac measurement package - Vascular measurement tools - Doppler velocity measurement - Automatic calculation packages - Annotation and reporting tools |  |  |  |  |  |
| - Workflow & Clinical Decision Support [MUST] |  |  |  |  |  |
| - Preset protocols for anesthesia/recovery applications - Quick exam workflow (<3 steps preferred) - Freeze/store/compare image function - Split-screen comparison - Automatic exam labeling |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Medical-grade display ≥ 15 inches. - High-resolution monitor - Tiltable and rotatable display - Touchscreen or programmable control panel - Backlit keyboard - User customizable presets |  |  |  |  |  |
| - Data Management & Connectivity [MUST] |  |  |  |  |  |
| - Internal storage for patient studies - DICOM compatibility - LAN connectivity - USB export capability - PACS integration capability - Network printing support |  |  |  |  |  |
| - Mobility & Ergonomics [MUST] |  |  |  |  |  |
| - Medical-grade cart with lockable wheels - Compact footprint suitable for recovery room - Cable management system - Height-adjustable control panel |  |  |  |  |  |
| - Power & Operational Reliability [MUST] |  |  |  |  |  |
| - AC 100–240 V - Internal battery operation ≥ 30 minutes - Instant resume after standby - Continuous operation without overheating |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Ultrasound main system – 1 - Phased array probe – 1 - Convex probe – 1 - Linear probe – 1 - Gel holders and probe holders - Footswitch (preferred) - User and service manuals |  |  |  |  |  |
| - Installation & Acceptance Testing [MUST] |  |  |  |  |  |
| - Acceptance testing shall include: - Image quality verification - Doppler function verification - Probe functionality testing - Data storage and export testing - Clinical demonstration in recovery setting - Acceptance checklist signed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - Clinical application training for anesthesia and recovery staff - Advanced POCUS workflow training - Biomedical maintenance training |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 3-year comprehensive warranty - Preventive maintenance included - Software updates included - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-37 (Diagnostic Ultrasound Equipment) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/