MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 45

| Priority | 2 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | HD Flexible Video Scope | Installation Location | Operation Room 2 |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The HD flexible video scope system shall provide high-definition endoscopic visualization for diagnostic and therapeutic procedures requiring flexible endoscopic access. - The system shall support safe navigation, precise visualization of anatomical structures, and compatibility with minimally invasive surgical workflows. - The system shall provide diagnostic as well as therapeutic interventions for gastrointestinal as well as ENT, pulmonology disciplines. Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade high-definition flexible video endoscope - Integrated digital imaging system (video processor compatible) - Designed for repeated clinical sterilization/disinfection cycles - Ergonomic control section suitable for prolonged procedures - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Imaging Performance [MUST] |  |  |  |  |  |
| - Processor must have 1920x1080p resolution with auto white balance, color correction, picture in picture mode. - The processor must have image enhancement (digital or optical) for tissue differentiation, detection and characterization - The processor must have inbuilt LED light source - The processor shall have inbuilt air and water feed system to the endoscope |  |  |  |  |  |
| - Optical & Illumination System [MUST] |  |  |  |  |  |
| - High-intensity LED or equivalent light source compatibility - Uniform illumination without peripheral darkening - Automatic brightness adjustment - Anti-fog capability or rapid fog recovery |  |  |  |  |  |
| - Insertion Tube & Mechanical Performance [MUST] |  |  |  |  |  |
| - Endoscope Upper GI: - Image sensor: 1080p CMOS technology with autofocus - Should have integrated waterjet function - Working length: Minimum 1050mm - Insertion tube diameter: Minimum 9.5mm - Distal end diameter: Maximum 11mm - Field of view: 140 - Depth of Field: Up/Down/Left/Right (210/120/120/120) - Endoscope Lower GI - Image sensor: 1080p CMOS technology with autofocus - Should have integrated waterjet function - Working length: Minimum 1700mm - Insertion tube diameter: 11-12mm - Distal end diameter: Maximum 12mm - Field of view: 140 - Depth of Field: Up/Down/Left/Right (180/180/160/160) |  |  |  |  |  |
| - Working Channel [MUST] |  |  |  |  |  |
| - Instrument channel allowing suction and accessory insertion - Instrument channel: Minimum 3.2mm for Upper GI endoscope and minimum 3.5mm for Lower GI endoscope - Leak-proof design - Easy flushing and cleaning pathway |  |  |  |  |  |
| - Control Section Ergonomics [MUST] |  |  |  |  |  |
| - One-hand operable angulation controls - Programmable function buttons - Suction and irrigation control - Non-slip medical-grade grip |  |  |  |  |  |
| - Image Processing & Clinical Enhancement [MUST] |  |  |  |  |  |
| - System shall support advanced visualization functions: - Edge enhancement - Contrast enhancement - Digital chromo-endoscopy or equivalent technology (acceptable) - Freeze and image capture - Video recording capability |  |  |  |  |  |
| - Video Processor & Connectivity [MUST] |  |  |  |  |  |
| - Compatible HD video processor included or specified - HDMI/DVI or equivalent HD output - Medical-grade monitor compatibility - USB or network image storage - Recording integration capability with inbuilt HDD |  |  |  |  |  |
| - Safety & Reliability [MUST] |  |  |  |  |  |
| - Electrical isolation compliant with medical standards - Automatic system self-check - Overheat protection - Leak testing capability prior to use |  |  |  |  |  |
| - Reprocessing & Infection Control [MUST] |  |  |  |  |  |
| - Fully immersible design - Compatible with standard high-level disinfection processes - Smooth surfaces minimizing biofilm retention - Manufacturer-approved reprocessing protocol provided - The system shall support leak testing prior to cleaning and disinfection procedures. |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each system shall include: - HD flexible video scope (upper GI) – 1 - HD flexible video scope (lower GI) – 1 - Compatible video processor – 1 - Light source (LED or equivalent) – 1 - Medical grade HD monitor compatibility set – 1 - Leak tester – 1 - Cleaning adapters – complete set - Starter biopsy/working accessories (basic set) - User & service manuals |  |  |  |  |  |
| - Installation & Acceptance Testing [MUST] |  |  |  |  |  |
| - Image quality verification - Articulation performance test - Leak integrity test - Illumination uniformity verification - Recording and storage function test - Acceptance checklist signed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - Clinical operation training - Reprocessing and infection control training - Biomedical maintenance training |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 3-year comprehensive warranty - Preventive maintenance included - Response time ≤ 48 hours - Spare parts availability ≥ 7–10 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-18 (Endoscopic equipment safety) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/