MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 46

| Priority | 2 | 3 | Floor | 1F-OT | Quantity | 1EA | 1EA |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  | Surgical Head Light | Installation Location | Operation Room 1 |  |  |
|  |  |  |  |  | Operation Room 5 |  |  |
| Specification Level |  |  | Low | Remarks |  |  |  |
| General Specification |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |
| - The headlight system shall provide focused, shadow-reduced illumination aligned with the surgeon’s line of sight for surgical and minor procedures, supporting accurate tissue visualization with minimal heat generation and improved operator comfort. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |
| - LED surgical headlight - Lightweight ergonomic headband - Adjustable and stable fixation - Continuous operation suitable for long procedures - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |
| - Illumination Performance [MUST] |  |  |  |  |  |  |  |
| - High-intensity focused illumination suitable for surgical visualization (manufacturer standard) - Adjustable brightness levels - Spot size adjustable from approx. 20–80 mm - Uniform light distribution without hot spots - Color temperature: 4,000–6,000 K - CRI ≥ 90 |  |  |  |  |  |  |  |
| - Optical & Thermal Safety [MUST] |  |  |  |  |  |  |  |
| - Low heat emission to patient and operator - No UV/IR harmful radiation - Silent cooling system |  |  |  |  |  |  |  |
| - Mechanical & Ergonomics [MUST] |  |  |  |  |  |  |  |
| - Lightweight head-mounted design suitable for prolonged clinical use (approximately ≤350 g preferred). - Balanced design for long-term comfort - Flexible beam positioning - Stable locking mechanism |  |  |  |  |  |  |  |
| - Power System [MUST] |  |  |  |  |  |  |  |
| - Rechargeable battery pack - Continuous runtime ≥ 6 hours - Charging time ≤ 3 hours - Battery level indicator - Minimum two batteries supplied for uninterrupted use - Battery system shall support continuous clinical operation with replaceable or hot-swappable battery capability acceptable. |  |  |  |  |  |  |  |
| - Durability [MUST] |  |  |  |  |  |  |  |
| - Resistant to routine disinfection - Cable strain relief protection - Drop-resistant construction |  |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |  |
| - Headlight – 1 - Battery pack – 2 - Charger – 1 - Waist belt/clip – 1 - Carry case – 1 - User manual – 1 |  |  |  |  |  |  |  |
| - Warranty & Service [MUST] |  |  |  |  |  |  |  |
| - Minimum 2-year warranty - Spare parts availability ≥ 5 years |  |  |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance), if applicable - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |

/End/