MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 47

| Priority | 2 | 3 | Floor | 1F-OT | Quantity | 1EA | 1EA |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  | Irrigation Pump | Installation Location | Operation Room 3 |  |  |
|  |  |  |  |  | Operation Room 4 |  |  |
| Specification Level |  |  | Mid | Remarks |  |  |  |
| General Specification |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |
| - The system shall provide controlled irrigation and suction during endoscopic and fluid-assisted surgical procedures to maintain clear visualization and efficient fluid removal. The system shall provide stable and controllable irrigation pressure and flow suitable for clinical use in urology, arthroscopy, and general surgery. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |
| - Integrated irrigation and suction control system - Digital pressure/flow control - Compact mobile design - Quiet operation - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |
| - Irrigation Performance [MUST] |  |  |  |  |  |  |  |
| - Adjustable irrigation flow suitable for endoscopic surgical procedures (approximately up to 1000 mL/min or equivalent). - Pressure control range: approx. 0–300 mmHg - Stable continuous flow without pulsation - Digital pressure display |  |  |  |  |  |  |  |
| - Suction Performance [MUST] |  |  |  |  |  |  |  |
| - Suction capability suitable for controlled surgical fluid evacuation and compatible with standard operating room suction systems. - Adjustable suction levels - Rapid fluid evacuation capability - Overflow and backflow protection |  |  |  |  |  |  |  |
| - Control & Interface [MUST] |  |  |  |  |  |  |  |
| - Digital display for pressure/flow - Preset programs - Footswitch activation (dual pedal) - Audible/visual alarms |  |  |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |  |  |
| - Over-pressure protection - Fluid level detection - Automatic shutdown on abnormal pressure - Electrical leakage protection - Automatic detection and alarm for tubing blockage or abnormal flow conditions |  |  |  |  |  |  |  |
| - Noise Level [MUST] |  |  |  |  |  |  |  |
| - Noise level appropriate for operating room use (approximately ≤65 dB preferred) |  |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |  |
| - Main unit – 1 - Footswitch – 1 - Irrigation tubing sets – minimum 10 - Suction tubing sets – minimum 10 - Collection bottles (2–3 L capacity) – minimum 2 - Pole or trolley – 1 - Power cable – 1 - Manuals – 1 set |  |  |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |  |  |
| - On-site installation and testing - User training for OR staff |  |  |  |  |  |  |  |
| - Warranty & Service [MUST] |  |  |  |  |  |  |  |
| - Minimum 3-year warranty - Preventive maintenance included - Response ≤ 48 hours - Spare parts ≥ 7 years |  |  |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |

/End/