MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 48

| Priority | 2 | Floor | 3F(Ward) | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Bladder Scanner | Installation Location | Treatment Room |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The bladder scanner shall provide non-invasive ultrasound measurement of urinary bladder volume for assessment of urinary retention and post-void residual urine (PVR). - The system shall support rapid bedside scanning with guided scanning operations to minimize operator dependency and support reduction of unnecessary catheterization. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Portable bladder ultrasound scanner with automatic bladder volume calculation - Real-time image guidance and scan quality indicator to support non-expert users - Suitable for adult and pediatric patients - Lightweight handheld probe with ergonomic grip - Disinfectant-resistant housing and probe surface - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Measurement & Imaging Performance [MUST] |  |  |  |  |  |
| - 3D or multi-plane scanning method for automatic volume estimation - Measurement range: 0–999 mL (minimum) - Measurement accuracy approximately ±15% or ±15 mL (whichever greater), within manufacturer specified measurement range - Repeatability: ±5% or better (manufacturer standard) - Scan time per measurement ≤ 5 seconds - Auto-identification of bladder boundary with on-screen bladder outline display - Capability to perform pre-void and post-void measurement and comparison - Automatic bladder volume calculation following ultrasound scanning |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Color display size ≥ 7 inches - Numeric display of bladder volume (mL) - Real-time bladder image with aiming/target guidance - Patient type selection: adult / pediatric - Patient ID entry and exam timestamp - Exam history review on device - On-screen scanning guidance or aiming assistance to support correct probe positioning shall be provided |  |  |  |  |  |
| - Data Storage & Connectivity [MUST] |  |  |  |  |  |
| - Internal storage capacity: minimum 500 examinations - USB export of reports (PDF/image) or equivalent - Optional printer connectivity acceptable - Date/time setting and retention during power off |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - Rechargeable battery operation - Continuous use ≥ 6 hours - Charging dock or charger included - Battery level indicator and low-battery alarm |  |  |  |  |  |
| - Probe & Durability [MUST] |  |  |  |  |  |
| - Probe cable strain relief and durable connector design - Waterproof probe head suitable for gel-based scanning and cleaning - Device protection: drop resistance suitable for bedside use (manufacturer specification) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Main console/unit – 1 - Bladder scanning probe – 1 - Charging dock/charger – 1 - Carry case – 1 - Ultrasound gel (starter supply) – minimum 2 bottles - Probe cover (starter pack) – minimum 50 pcs (if applicable) - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - On-site commissioning and functional testing - Verification of volume measurement function and data export - Acceptance checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for nurses and clinicians - Training includes scanning technique, PVR workflow, cleaning, and data export |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-37 or equivalent (ultrasound) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/