MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 49

| Priority | 3 | Floor | GF | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | High-Frequency Electrosurgical Unit | Installation Location | Obstetrics and Gynecology Treatment Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The electrosurgical unit shall provide safe and precise tissue cutting and coagulation during gynecological and minor surgical procedures. - The system shall support monopolar and bipolar operation with stable output regulation and patient return electrode safety monitoring. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled HF electrosurgical generator - Monopolar and bipolar operation - OR-grade continuous operation - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Output Modes [MUST] |  |  |  |  |  |
| - Monopolar: - Pure cut - Blend cut - Forced coagulation - Spray coagulation or equivalent coagulation mode acceptable - Soft coagulation or equivalent low-voltage coag mode - Bipolar: - Bipolar coagulation - Micro/precise bipolar mode |  |  |  |  |  |
| - Output Performance [MUST] |  |  |  |  |  |
| - Monopolar cut maximum power ≥ 300 W or equivalent suitable for gynecological surgical procedures. - Monopolar coag maximum power ≥ 120 W - Bipolar maximum power ≥ 70 W - Automatic output regulation with tissue impedance compensation - Fine power adjustment and stable output across operating range |  |  |  |  |  |
| - Patient Safety & Monitoring [MUST] |  |  |  |  |  |
| - Return electrode contact quality monitoring (REM/CQM) - Audible/visual alarm and automatic output stop on poor pad contact - Isolated output circuits - Self-test at startup - Overload, short-circuit, and over-temperature protection - Activation indicators for monopolar/bipolar outputs - Activation indicators for monopolar/bipolar outputs |  |  |  |  |  |
| - User Interface & Control [MUST] |  |  |  |  |  |
| - Digital display for mode and power - Preset program memory (minimum 10 programs) - Footswitch and handpiece activation supported - Clear mode selection and activation status display |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate OR use: - ESU main generator – 1 - Monopolar Dual footswitch – 1 - Monopolar handpiece/pencils – minimum 10 pcs - Monopolar cable – 1 - Bipolar forceps – 2 pcs - Bipolar cable – 1 - Bipolar footswitch- 1 - Patient return electrodes (disposable) – minimum 50 pcs - Return electrode cable – 1 - Trolley/cart – 1 - Power cable – 1 - User & service manuals – 1 set |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Vessel sealing module and handpieces - Smoke evacuation interface/module - Additional bipolar instruments |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - On-site commissioning and verification of all modes - Electrical safety test verification - REM/CQM function test - Acceptance checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for surgeons, nurses, and biomedical technicians - Includes safe pad placement and alarm management |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-2 (HF surgical equipment) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/