MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.5

| Priority | 1 | 2 | Floor | 1F-DR | GF | Quantity | 2EA | 1EA |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  |  | Fetal Monitor | Installation Location |  | Labor Room |  |  |
|  |  |  |  |  |  | Obstetrics & Gynecology consultation room |  |  |
| Specification Level |  |  | Mid | Remarks |  |  |  |  |
| General Specification |  |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |  |  |
| - The fetal monitor shall be used for antepartum and intrapartum monitoring of fetal heart rate (FHR) and uterine activity (TOCO) during pregnancy and labor. - The system shall provide reliable continuous monitoring suitable for routine and emergency obstetric care. - Equivalent or higher performance systems are acceptable. Specifications shall be vendor-neutral and performance-based. |  |  |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |  |  |
| - Compact bedside fetal monitor (CTG) - User-friendly operation suitable for nurses and midwives - Durable construction for daily clinical use - Easy cleaning and disinfectant-resistant surfaces - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |  |  |
| - Monitoring Functions [MUST] |  |  |  |  |  |  |  |  |
| - Fetal Monitoring - Ultrasound Doppler FHR monitoring for single fetus - Measurement range appropriate for obstetric use (approx. 50–240 bpm) - Cross-channel interference prevention or equivalent signal protection - Uterine Activity - External TOCO monitoring - Maternal Monitoring - Maternal heart rate or SpO₂ capability (via integrated or shared module) - Alarm & Safety - Audible and visual alarms - Adjustable alarm limits |  |  |  |  |  |  |  |  |
| - Display & Recording [MUST] |  |  |  |  |  |  |  |  |
| - Color display ≥ 6 inches - Clear waveform and numeric display - Integrated thermal printer/recorder - Selectable paper speeds suitable for CTG recording - Internal memory or event storage capability |  |  |  |  |  |  |  |  |
| - Connectivity & Data [MUST] |  |  |  |  |  |  |  |  |
| - USB or LAN connectivity for data export - Capability for integration with hospital information system or obstetric information system - Standard data formats |  |  |  |  |  |  |  |  |
| - Power & Portability [MUST] |  |  |  |  |  |  |  |  |
| - AC power operation - Internal rechargeable battery backup ≥ 2 hours minimum - Lightweight portable design for room-to-room movement |  |  |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |  |  |
| All components required for full clinical operation must be included. - Each unit shall include: - Main monitor unit - Ultrasound transducer × 2 - TOCO transducer × 1 - Transducer belts/straps – complete set - Event marker button - Printer paper – starter pack - Power cable - All required software and licenses - User and service manuals |  |  |  |  |  |  |  |  |
| - Optional Features (price separately) |  |  |  |  |  |  |  |  |
| The following shall NOT be included in the base price and must be quoted separately: - Twin/triplet monitoring modules - Wireless/cordless transducers - Central monitoring station - Additional transducers or accessories |  |  |  |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |  |  |  |
| - Delivery, installation, and commissioning by qualified personnel - Functional testing on-site - Acceptance testing included |  |  |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |  |  |
| - On-site user training for clinicians and nurses - Training on monitoring, interpretation basics, cleaning, and routine checks |  |  |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, probes, and accessories - Response time ≤ 48 hours - Preventive maintenance during warranty included - Spare parts availability ≥ 7 years |  |  |  |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |  |  |

/END/