MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 50

| Priority | 3 | Floor | GF | Quantity | 2EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Side Lamp | Installation Location | Obstetrics and Gynecology Treatment Room, Orthopedics Treatment Room |  |
| Specification Level |  | Low | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The examination lamp shall provide localized clinical illumination for examination and minor procedures including suturing and dressing. - The lamp shall support stable illumination, accurate color rendering, and flexible positioning without drift during clinical use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - LED examination/procedure lamp - Floor-standing mobile type with lockable casters - Disinfectant-resistant surfaces and easy cleaning - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Illumination Performance [MUST] |  |  |  |  |  |
| - Illuminance ≥ 40,000 lux at 1 meter (minimum) - Adjustable brightness with at least 3 levels - Spot diameter adjustable: approx. 100–200 mm at 1 meter - Uniform illumination with shadow reduction - Color temperature: 4,000–5,000 K - CRI ≥ 90 |  |  |  |  |  |
| - Mechanical Performance [MUST] |  |  |  |  |  |
| - Multi-joint articulated arm with reach ≥ 800 mm - 360° rotation capability - Stable positioning without drift after adjustment - Height adjustable suitable for seated/bedside procedures - Sterilizable handle provided for repositioning |  |  |  |  |  |
| - Thermal & Noise [MUST] |  |  |  |  |  |
| - Low heat emission at working distance - Silent operation without fan noise impacting patient comfort |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Battery backup capability acceptable to support temporary operation during patient movement or power interruption. |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Lamp head – 1 - Articulated arm – 1 - Mobile base with casters and brakes – 1 - Sterilizable handle – minimum 2 - Power cable – 1 - User manual – 1 |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - Assembly and functional test on-site - Verification of illuminance and arm stability - Acceptance checklist completed |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year warranty - Spare parts availability ≥ 5 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance), if applicable - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/