MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 51

| Priority | 3 | Floor | 1F-OT, 2F(ward) 3F(ward) |  | Quantity | 2EA for each Treatment room, 1EA for NICU, Labor room |
| --- | --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Pulse Oximeter |  | Installation Location | Treatment Room(2F), Treatment Room(3F), Labor room, NICU |  |
| Specification Level |  | Low |  | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |  |
| - The pulse oximeter shall provide continuous non-invasive monitoring of arterial oxygen saturation (SpO₂) and pulse rate for adult, pediatric, and neonatal patients. - The system shall be suitable for bedside monitoring and spot-check monitoring in ward, emergency, critical care, and neonatal care settings. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |  |
| - Medical-grade bedside pulse oximeter (NOT fingertip consumer type) - Portable tabletop or handheld type with stand/pole mount - Continuous monitoring capability - Durable disinfectant-resistant housing - Easy-to-use interface for nursing staff - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |  |
| - Measurement Performance [MUST] |  |  |  |  |  |  |
| - SpO₂: - Measurement range: 0–100% - Accuracy: ±2% (70–100%) or better - Motion-tolerant performance suitable for restless patients - Low perfusion accuracy capability (manufacturer standard) - Suitable for neonatal and low-perfusion monitoring environments - Pulse Rate: - Range: minimum 25–240 bpm or wider - Accuracy: ±2 bpm or ±2% (whichever greater) - Perfusion: - Perfusion index (PI) display or equivalent acceptable |  |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |  |
| - Color display ≥ 3.5 inches - Continuous waveform (plethysmograph) display - Numeric display of SpO₂ and pulse rate - Adjustable brightness - Simple keypad or touch interface - Battery and alarm status indicators |  |  |  |  |  |  |
| - Alarm & Safety Functions [MUST] |  |  |  |  |  |  |
| - Audible and visual alarms - Adjustable alarm limits for SpO₂ and pulse rate - Sensor disconnection alarm - Low battery alarm - High/low priority alarm differentiation - Alarm silence with automatic reactivation |  |  |  |  |  |  |
| - Trend & Data Storage [MUST] |  |  |  |  |  |  |
| - Trend storage ≥ 24 hours - Event/alarm history review - Date/time stamping - USB or equivalent data export capability |  |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Internal rechargeable battery - Battery backup ≥ 6–8 hours continuous monitoring - Battery level indicator |  |  |  |  |  |  |
| - Probe Compatibility [MUST] |  |  |  |  |  |  |
| - Reusable SpO₂ probes included: - Adult – 1 - Pediatric – 1 - Neonatal – 1 - Compatible with commonly available standard disposable probes without mandatory proprietary consumables - Durable cable with strain relief |  |  |  |  |  |  |
| - Mounting & Portability [MUST] |  |  |  |  |  |  |
| - Tabletop or handheld configuration - IV pole mount or stand included - Lightweight and portable for bedside use |  |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Pulse oximeter main unit – 1 - Adult reusable probe – 1 - Pediatric reusable probe – 1 - Neonatal reusable probe – 1 - Pole clamp/stand – 1 - Rechargeable battery (installed) - Power adapter/cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |  |
| - Functional testing of measurement and alarms - Verification of probe operation - Acceptance checklist completed on-site |  |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |  |
| - On-site user training for nursing staff - Training includes alarm setting, probe care, and cleaning |  |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included - Spare parts availability ≥ 7 years |  |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-49 or equivalent (multiparameter/pulse oximetry monitoring) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |  |

/End/