MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 52

| Priority | 3 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Vacuum Delivery System | Installation Location | Delivery Room |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The vacuum delivery system shall be used in labor and delivery suites for assisted vaginal delivery procedures. - The system shall provide controlled and safe negative pressure for obstetric use to support operative vaginal delivery procedures and maternal and fetal safety. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electrically powered vacuum system suitable for obstetric use - Stable and adjustable suction control - Portable or trolley-mounted design - Durable medical-grade housing resistant to disinfectants - Quiet operation suitable for labor room environment - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Performance Specifications [MUST] |  |  |  |  |  |
| - Adjustable negative pressure range suitable for obstetric vacuum procedures (approximately up to –80 kPa or equivalent) - Fine pressure regulation with continuous adjustment - Vacuum build-up time suitable for clinical obstetric use (manufacturer standard) - Flow rate ≥ 20 L/min - Accurate vacuum gauge displaying kPa and mmHg - Automatic pressure release mechanism - Overflow protection system with float valve - Integrated hydrophobic bacterial/viral filter |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Over-vacuum protection - Audible/visual alarm for pressure deviation - Anti-backflow mechanism - Secure suction tubing connectors - Electrical safety compliant with IEC 60601-1 |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Dual power operation preferred (AC + internal rechargeable battery) - Battery runtime ≥ 30 minutes (if battery equipped) |  |  |  |  |  |
| - Noise & Mechanical Performance [MUST] |  |  |  |  |  |
| - Noise level ≤ 55 dB - Stable base or locking wheels for trolley version - Weight ≤ 10 kg (portable configuration) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Main vacuum pump unit – 1 - Suction regulator – 1 - Collection jar 1–1.5 L autoclavable – 2 - Silicone tubing sets – 2 - Hydrophobic bacterial filters – 10 - Stainless steel traction handle – 2 - Silicone vacuum cups (50 mm & 60 mm) – 1 each - Stainless steel bird cups (40/50/60 mm) – 1 each - Power cable – 1 - User & service manuals – 1 set - Calibration certificate – 1 |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |
| - On-site installation and calibration - Minimum 2 training sessions for OB/GYN staff |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year comprehensive warranty - Local authorized service agent in Nepal mandatory - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/