MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 53

| Priority | 3 | Floor | 1F-OT, 1F-DR | Quantity | 5 EA (1EA For each Operating Rooms, Recovery Room, Labor Room) |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Medical Blanket Warmer |  | Installation Location | Operating Room 1, Operating Room 2, Operating Room 3, Recovery Room Labor Room |
| Specification Level |  | Mid |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The blanket warmer shall provide controlled warming of surgical and recovery blankets to support maintenance of patient normothermia and thermal comfort. - The system shall be suitable for continuous daily clinical use in operating, recovery, and labor environments. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled warming cabinet - Fan-assisted convection heating for uniform temperature - Transparent door for blanket visibility - Durable stainless steel or medical-grade interior - Disinfectant-resistant external housing - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Temperature Performance [MUST] |  |  |  |  |  |
| - Adjustable temperature settings approximately within 32°C–43°C or wider clinically suitable range - Temperature accuracy ±1°C or better - Rapid warm-up to 38°C within 30 minutes or less - Uniform internal temperature distribution |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Independent over-temperature safety cutoff - Audible/visual alarm for overheating - Thermal sensor fault alarm - Auto shut-off in case of malfunction |  |  |  |  |  |
| - Capacity & Mechanical Performance [MUST] |  |  |  |  |  |
| - Capacity suitable for OR blanket demand - Internal capacity ≥ 150 L (minimum) - Door with magnetic seal - Lockable wheels (if mobile) - Noise ≤ 50 dB |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 220–240 V, 50/60 Hz - Stable power consumption suitable for hospital use |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Blanket warming cabinet – 1 - Internal shelves – 2 or more - Temperature control panel – integrated - Power cable – 1 - User & service manuals – 1 set - Calibration certificate – 1 |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |
| - On-site installation and testing - Minimum 2 training sessions for OR staff |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 2-year comprehensive warranty - Local service support required - Spare parts ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/