MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 54

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Intubation Fiberscope |  | Installation Location | Recovery Room |
| Specification Level |  | High |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The intubation scope system shall provide real-time visualization of airway anatomy to support safe endotracheal intubation, particularly in difficult airway situations. - The system shall enable controlled placement of endotracheal tubes under direct visualization to support reduction of airway trauma and hypoxic complications. - The system shall be suitable for anesthesia induction, emergency airway management, ICU airway intervention, and awake fiberoptic intubation procedures. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade flexible video intubation fiberscope - Designed specifically for airway management - Portable system allowing rapid deployment - Fast startup suitable for emergency use - Lightweight ergonomic design for single operator use - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. - The insertion tube and patient-contact components shall be compatible with high-level disinfection (HLD) or equivalent validated reprocessing methods in accordance with manufacturer instructions. |  |  |  |  |  |
| - Imaging Performance [MUST] |  |  |  |  |  |
| - High-resolution digital imaging (HD or equivalent clinical resolution) - CMOS imaging sensor at distal tip - Field of view ≥ 90° - Depth of field optimized for airway structures - Automatic exposure and brightness control - Stable image without flicker during movement - System shall provide clear visualization of vocal cords and tracheal anatomy. |  |  |  |  |  |
| - Insertion Tube & Mechanical Characteristics [MUST] |  |  |  |  |  |
| - Flexible insertion tube with atraumatic distal tip - Outer diameter suitable for adult airway use (approx. 4–6 mm range or equivalent) - Working length appropriate for adult intubation - Smooth articulation mechanism - Angulation performance: - Up ≥ 120° - Down ≥ 120° (or equivalent clinical maneuverability - High durability against repeated bending cycles - Waterproof construction for medical reprocessing |  |  |  |  |  |
| - Working Channel [MUST] |  |  |  |  |  |
| - Integrated working/suction channel - Compatible with oxygen insufflation and suction - Channel diameter adequate for secretion clearance - Easy flushing and cleaning pathway |  |  |  |  |  |
| - Illumination System [MUST] |  |  |  |  |  |
| - Integrated LED light source or equivalent illumination system acceptable - Uniform illumination without hotspot glare - Low heat emission at distal tip - Automatic brightness adjustment |  |  |  |  |  |
| - Clinical Safety Features [MUST] |  |  |  |  |  |
| - Rapid visual acquisition for emergency airway use - Anti-fog capability or rapid fog recovery - Smooth atraumatic tip design - Stable visualization during patient movement - Electrical isolation compliant with medical safety standards - The system shall support leak testing capability to verify scope integrity prior to cleaning and disinfection. |  |  |  |  |  |
| - Display & Control Interface [MUST] |  |  |  |  |  |
| - Portable monitor ≥ 7–10 inches or compatible HD display - Real-time video display with minimal latency - Image freeze and capture function - Video recording capability preferred - One-touch recording or snapshot control |  |  |  |  |  |
| - Operational Reliability [MUST] |  |  |  |  |  |
| - Instant-on or rapid boot capability - Continuous operation without overheating - Signal stability during prolonged procedures - Battery-powered operation ≥ 1 hour for emergency use |  |  |  |  |  |
| - Infection Control & Reprocessing [MUST] |  |  |  |  |  |
| - Fully immersible design - Compatible with high-level disinfection procedures - Smooth external surface minimizing contamination retention - Manufacturer-approved cleaning protocol provided - Leak testing capability prior to use |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each system shall include: - Video intubation fiberscope – 1 - Portable monitor/control unit – 1 - Charging system or power supply – 1 - Suction/oxygen adapters – complete set - Leak tester – 1 - Cleaning and disinfection adapters – complete set - Protective carrying case – 1 - User and service manuals – 1 set |  |  |  |  |  |
| - Installation & Acceptance Testing [MUST] |  |  |  |  |  |
| - Acceptance testing shall include: - Image clarity verification of airway structures - Articulation movement test - Illumination uniformity verification - Leak integrity test - Recording and display functionality test - Acceptance checklist signed |  |  |  |  |  |
| - Training [MUST |  |  |  |  |  |
| - Clinical airway management training - Cleaning and reprocessing training - Biomedical maintenance orientation |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 3-year comprehensive warranty - Preventive maintenance included - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/