MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 55

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | High Flow Blood & Fluid Warmer |  | Installation Location | Recovery Room |
| Specification Level |  | Mid |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The fluid warmer shall warm blood and intravenous fluids during transfusion and resuscitation to prevent hypothermia. - The system shall deliver warmed fluids with controlled temperature and flow suitable for emergency resuscitation and surgical use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Compact portable warmer - Rapid heating technology (dry heat or equivalent) - Suitable for OR and emergency use - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Temperature Performance [MUST] |  |  |  |  |  |
| - Adjustable outlet temperature approximately within 37–41°C suitable for blood and fluid warming - Temperature accuracy: ±1°C or better - Rapid heating from room temperature - Stable outlet temperature at high flow |  |  |  |  |  |
| - Flow Performance [MUST] |  |  |  |  |  |
| - Flow rate capability suitable for high-flow fluid and blood warming during transfusion and resuscitation (approximately up to 1000 mL/min or equivalent) - Suitable for rapid transfusion - No significant flow resistance |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Over-temperature protection - Air detection - Dry heat protection - Audible/visual alarms - Automatic shutdown on fault |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |
| - Digital temperature display - Alarm indicators - Simple controls |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V - Continuous operation |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Main warmer unit – 1 - Disposable warming sets – minimum 20 - Pole clamp/stand – 1 - Power cable – 1 - User manual – 1 |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |
| - Functional testing on-site - Staff training |  |  |  |  |  |
| - Warranty & Service [MUST] |  |  |  |  |  |
| - Minimum 3-year warranty - Preventive maintenance included - Response ≤ 48 hours - Spare parts ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/