MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 56

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Video Laryngoscope |  | Installation Location | Recovery Room |
| Specification Level |  | High |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The video laryngoscope system shall provide real-time visualization of airway anatomy to assist routine and difficult intubation procedures. - The system shall support visualization during intubation procedures and facilitate safe airway management. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade video laryngoscope (NOT consumer handheld type) - Portable, battery-operated system suitable for bedside and emergency use - Integrated or detachable color monitor - Durable housing resistant to routine hospital disinfectants - Designed for repeated clinical use - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Imaging Performance [MUST] |  |  |  |  |  |
| - Resolution: HD (minimum 1280 × 720) or higher - CMOS or CCD digital imaging sensor - Wide field of view ≥ 60° - Depth of field suitable for airway visualization (approx. 10–100 mm) - Anti-fog or rapid warm-up lens technology - Clear visualization without external light source - Real-time video without noticeable lag |  |  |  |  |  |
| - Display Requirements [MUST] |  |  |  |  |  |
| - Color display size ≥ 4 inches - Adjustable brightness - Anti-glare screen - Image rotation or orientation adjustment - Simultaneous live image and status indicators |  |  |  |  |  |
| - Blade System [MUST] |  |  |  |  |  |
| - Reusable or disposable blade options acceptable. - Blade configurations suitable for adult, pediatric, and neonatal airway management shall be provided. - Macintosh-type or equivalent clinically accepted blade geometries are acceptable. - Easy attachment and removal mechanism. - Autoclavable (if reusable) or sterile single-use. - Medical-grade material resistant to repeated cleaning and disinfection. |  |  |  |  |  |
| - Mechanical & Durability [MUST] |  |  |  |  |  |
| - Waterproof or splash-proof protection: IPX7 or equivalent - Drop resistance suitable for bedside use (manufacturer specification) - Robust cable or integrated handle design - Weight suitable for one-hand operation |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - Rechargeable lithium battery - Continuous runtime ≥ 4–6 hours - Charging time ≤ 3 hours - Replaceable or spare battery included - Battery level indicator |  |  |  |  |  |
| - Recording & Data Management [MUST] |  |  |  |  |  |
| - Image capture (still image) - Video recording capability - Internal storage ≥ 8 GB or removable SD/USB - USB export or equivalent data transfer |  |  |  |  |  |
| - Safety & Infection Control [MUST] |  |  |  |  |  |
| - Smooth surfaces for easy cleaning - Compatible with hospital disinfectants - No sharp edges - Electrical safety compliant with IEC 60601-1 |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Video laryngoscope handle – 1 - Monitor/display – 1 - Adult blades – minimum 2 sizes - Pediatric blade – 1 - Neonatal blade – 1 - Rechargeable batteries – minimum 2 - Charger or docking station – 1 - Carry case – 1 - USB cable – 1 - User & service manuals – 1 set |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - Functional verification of image quality and recording - Battery and charging test - Acceptance checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for ER/OR/ICU staff - Training includes blade attachment, cleaning, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 2 years - Covers parts, labor, and travel - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/