MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 57

| Priority | 3 | Floor | 1F-OT | Quantity | 2 EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Digital Storage System |  | Installation Location | Operating Room 1 Operating Room 2 |
| Specification Level |  | Mid |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The digital storage system shall capture, store, manage, and archive medical images and videos generated from surgical cameras, endoscopy systems, and imaging equipment. - The system shall support secure data storage, easy retrieval, and backup and data redundancy capability. - The system shall function as a medical image documentation and archiving system suitable for operating room clinical workflow and shall support integration with hospital PACS where applicable. - Equivalent or higher performance systems are acceptable. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - The system shall be designed and validated for medical image recording and clinical documentation use. - Medical-grade image and video recording system suitable for operating room applications. - Compatible with 4K, Full HD and HD surgical camera and endoscopy systems. - DICOM compatible for medical image management and integration. - Secure and stable storage architecture shall be provided. - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Recording Performance [MUST] |  |  |  |  |  |
| - Support 4K and must be compatible with lower resolutions (Full HD, HD) - Must support HDMI 2.0 or 12G-SDI interfaces to handle the high bandwidth required for 4K - Real-time image capture (still and video) - Simultaneous display and recording capability - Minimum recording frame rate 30 fps |  |  |  |  |  |
| - Storage Capacity [MUST] |  |  |  |  |  |
| - Internal storage ≥ 8 TB - Expandable storage capability - Automatic backup function (external HDD or network) - RAID redundancy configuration shall be provided for data protection - Must support H.265 (HEVC) video compression. It must also support H.264 (AVC) for compatibility with older playback software and lower-resolution files. |  |  |  |  |  |
| - Data Management & Security [MUST] |  |  |  |  |  |
| - Patient ID input and tagging - Date/time stamping - User access control with password protection - Audit log capability - USB export and network transfer - DICOM 3.0 compatibility - Role-based user access control with administrator privilege management |  |  |  |  |  |
| - Connectivity [MUST] |  |  |  |  |  |
| - LAN port - USB ports (minimum 2) - HDMI or 12G-SDI equivalent video input/output - Integration with hospital EMR or PACS capability - Capability for integration with hospital PACS and EMR systems via standard network protocols |  |  |  |  |  |
| - Display & Interface [MUST] |  |  |  |  |  |
| - User-friendly graphical interface - Search and retrieval function - Thumbnail preview - Multilingual support (English mandatory) |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V - Surge protection recommended |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Main storage unit – 1 - External backup HDD – 1 (minimum 8 TB) - Network cables – 1 set - User & service manuals – 1 set |  |  |  |  |  |
| - Installation & Training [MUST] |  |  |  |  |  |
| - On-site installation and network configuration - Integration testing with OR camera system - Staff training (minimum 2 sessions) |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Minimum 3-year warranty - Software updates included - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - CE marking or equivalent regulatory approval (if classified as medical device) - IEC 60601-1 (if applicable) - ISO 13485 or ISO 9001 certified manufacturer - IEC 62304 compliant software lifecycle process (if medical device software applicable) |  |  |  |  |  |

/End/