MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 58

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Endo-Irrigator |  | Installation Location | Operating Room 2 |
| Specification Level |  | Mid |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The endoscopic irrigation system shall provide controlled and continuous irrigation during minimally invasive surgical procedures including urology, gynecology, and general endoscopic procedures, and where applicable arthroscopic procedures. - The system shall support maintenance of clear visualization by delivering stable fluid flow and controllable irrigation pressure without sudden fluctuations. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled irrigation pump system - Digital pressure and flow control - Continuous-duty operation suitable for prolonged surgery - Compact OR-compatible configuration - Mobile trolley or pole mount - Disinfectant-resistant medical-grade housing - Must be capable of cleaning the tip of the endoscope - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Flow Performance [MUST] |  |  |  |  |  |
| - Adjustable flow rate suitable for continuous endoscopic irrigation (approximately up to 800–1000 mL/min or equivalent) - Flow accuracy: ±10% or better - Stable continuous flow without pulsation - Rapid response to flow adjustment - Flow maintained under surgical load |  |  |  |  |  |
| - Pressure Performance [MUST] |  |  |  |  |  |
| - Adjustable pressure range suitable for endoscopic irrigation procedures (approximately up to 200 mmHg or equivalent) - Pressure accuracy: ±10 mmHg or better - Stable pressure control without overshoot - Automatic compensation to maintain preset pressure - Digital pressure display |  |  |  |  |  |
| - Pump & Mechanical Performance [MUST] |  |  |  |  |  |
| - Roller pump, diaphragm pump, or equivalent surgical-grade pump - Continuous operation capability ≥ 60 minutes without overheating - Noise level ≤ 60 dB during operation - Smooth and vibration-minimized operation |  |  |  |  |  |
| - Fluid Handling System [MUST] |  |  |  |  |  |
| - Compatible with standard sterile saline bags - Container or bag holder capacity ≥ 3 liters - Quick-connect tubing system - Anti-backflow mechanism - Easy tubing replacement during surgery |  |  |  |  |  |
| - Control Interface [MUST] |  |  |  |  |  |
| - Digital display of flow and pressure - Preset memory programs (minimum 5) - Footswitch activation (dual pedal preferred) - Clear status indicators |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Over-pressure protection - Automatic shutdown in abnormal condition - Fluid level detection alarm - Audible and visual alarms |  |  |  |  |  |
| - Sterilization & Durability [MUST] |  |  |  |  |  |
| - Autoclavable reusable components (if applicable) - Tubing compatible with sterilization or disposable sterile sets - Designed to withstand ≥ 500 sterilization cycles (for reusable parts) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Main irrigation console – 1 - Mobile trolley or pole mount – 1 - Footswitch – 1 - Irrigation tubing sets – minimum 10 - Endoscope cleaning Sheath(10mm diameter)-1 - Fluid bag holders – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site installation and functional testing - Verification of flow and pressure accuracy - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for OR nurses and surgeons - Includes operation, tubing setup, sterilization, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel - Preventive maintenance included - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/