MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 59

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Surgical Drill System |  | Installation Location | Operating Room 1 |
| Specification Level |  | High |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The surgical drill system shall be used for Neurosurgery, - The system shall provide high torque, stable rotational speed, low vibration, and ergonomic handling designed to reduce operator fatigue, supporting precise surgical performance during prolonged procedures. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade electric surgical drill system - Continuous-duty operation suitable for operating room environment - Lightweight, ergonomic handpiece for prolonged use - Low vibration and low noise operation - Disinfectant-resistant and sterilizable components - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Performance Specifications [MUST] |  |  |  |  |  |
| - Speed: - Adjustable speed range (Maximum speed ≥ 70,000 RPM for Neuro applications, 30,000 RPM for ENT applications). - Variable speed control via trigger or footswitch - Stable rotation without speed fluctuation under load - Torque: - High torque under low-speed surgical operation - No stall or sudden stop during standard surgical load - Vibration & Noise: - Low vibration design for precise control - Noise level minimized for operating room use (manufacturer specification required) - Run Stability: - Smooth rotation without excessive heat generation - Continuous use capability ≥ 60 minutes without performance degradation |  |  |  |  |  |
| - Control System [MUST] |  |  |  |  |  |
| - Single console for all different motors and applications - Variable speed trigger control and/or footswitch - Forward and reverse rotation - Immediate stop function - Clear speed setting indicators - Integrated pump with dynamic flow control (up to 100 cc/min) for irrigation controlled by same footswitch |  |  |  |  |  |
| - Mechanical & Ergonomics [MUST] |  |  |  |  |  |
| - Handpiece weight ≤ 1.5 kg - Non-slip ergonomic grip - Balanced design to reduce surgeon fatigue - Quick-connect mechanism for drill bits and attachments - Secure locking to prevent accidental release |  |  |  |  |  |
| - Sterilization & Durability [MUST] |  |  |  |  |  |
| - Autoclavable handpiece and attachments - Compatible with ≥ 134°C steam sterilization - Designed to withstand ≥ 500 sterilization cycles without degradation - Waterproof or sealed design to prevent fluid ingress - Stainless steel or medical-grade alloy construction |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| (Choose one; bidder to specify) - For electric type: - AC 100–240 V operation - OR-grade power supply |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Overheat protection - Electrical leakage protection - Secure cable strain relief - Overload protection - Emergency stop capability |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate surgical use: - Main control unit – 1 - Motor for Neurosurgery - Perforator for Neurosurgery - Attachment for Neurosurgery (Straight Attachment 9cm – 1, Straight Attachment Long 14cm – 1, Angled Attachment Medium 10cm – 1, Footed Craniotome Driver – 1, Perforator Driver – 1) - Motor for Microdebrider – 1 - Footswitch – 1 - Chuck/collet system – 1 - Drill bits (starter set, assorted sizes) – minimum 10 - Reamer or burr attachments – minimum 5 - Sterilization tray/case – 1 - Spare seals/gaskets – 1 set - Power cable or charger – 1 - User & service manuals – 1 set |  |  |  |  |  |
| - Installation & Acceptance [MUST] |  |  |  |  |  |
| - Functional testing of speed, torque, and attachments - Electrical safety testing - Acceptance checklist completed on-site |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for surgeons, nurses, and biomedical engineers - Includes assembly, sterilization, and maintenance procedures |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel - Preventive maintenance included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/