MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.6

| Priority | 1 | Floor | 1F-DR | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Incubator (Portable) | Installation Location | Delivery Room |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The portable transport incubator shall be used to provide a stable, thermally controlled, and protected microenvironment for neonatal and premature infants during intra-hospital and inter-department transport. - The system shall provide controlled temperature, humidity, and oxygen support to promote thermal stability and physiological protection during transport. - The equipment shall be suitable for continuous clinical use and safe patient transport within the hospital. - The system shall be intended for short-duration transport stabilization rather than long-term intensive care use. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade neonatal transport incubator (NOT basic warming box) - Fully enclosed double-wall acrylic glass hood design for thermal stability - Compact and lightweight structure suitable for transport - Shock-resistant construction for safe movement - Smooth surfaces resistant to hospital disinfectants - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Temperature Control Performance [MUST] |  |  |  |  |  |
| - Air temperature control range: 20–38°C or wider - Temperature accuracy: ±0.2°C or better - Temperature uniformity within chamber: less than ± 1.5°C or better - Servo-controlled skin temperature mode with patient probe - Air temperature mode selectable - Warm-up time to 36°C ≤ 20 minutes - Automatic compensation for door opening or heat loss |  |  |  |  |  |
| - Humidity Control [MUST] |  |  |  |  |  |
| - Relative humidity control capability ≥ 50–80% (if integrated humidification) - Stable humidity maintenance - Easy refill water chamber - Condensation control design |  |  |  |  |  |
| - Accessories Included [MUST] |  |  |  |  |  |
| - Forced air circulation system for uniform heat distribution - Low airflow noise and minimal drafts - Air filtration system (bacterial filter or equivalent) - Noise level ≤ 52 dB during operation |  |  |  |  |  |
| - Oxygen Support Capability [MUST] |  |  |  |  |  |
| - Oxygen inlet port for supplemental oxygen - Oxygen concentration monitoring capability. - Safe and controlled oxygen delivery without excessive heat loss |  |  |  |  |  |
| - Hood & Patient Compartment [MUST] |  |  |  |  |  |
| - Transparent double-wall hood for clear visibility - Multiple access ports with soft sleeves - Quick-open side doors for emergency access - Mattress size suitable for neonatal use (minimum 60 × 30 cm approx.) - Radiolucent mattress base for X-ray without patient removal - Easy removal and cleaning |  |  |  |  |  |
| - Transport & Mechanical Design [MUST] |  |  |  |  |  |
| - Integrated trolley or transport cart with lockable wheels - Secure fastening system for ambulance/bedside use (if applicable) - Compartment of integrated two oxygen cylinder holder of 4 liter - Anti-vibration/shock absorption design - Total weight suitable for safe handling - Battery compartment protected from impact |  |  |  |  |  |
| - Power System [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz operation - Internal rechargeable battery - Battery runtime ≥ 3 hours continuous operation - Battery charging time ≤ 4 hours - Battery level indicator and low battery alarm |  |  |  |  |  |
| - Monitoring & Display [MUST] |  |  |  |  |  |
| - 5 inch or more color LCD touch Digital display of: - Air temperature - Skin temperature - Set temperature - Battery status - Clear alarm indicators - Easy-to-use control interface |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - High/low temperature alarm - Sensor failure alarm - Power failure alarm - Battery low alarm - Door open alarm - Overheat protection |  |  |  |  |  |
| - Continuous Operation Capability [MUST] |  |  |  |  |  |
| - Designed for continuous operation ≥ 4 hours without performance degradation - Stable thermal environment during movement |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all components necessary for immediate clinical use: - Portable incubator main unit – 1 - Transport trolley/cart – 1 - 2 units of 4 liter oxygen cylinder - Mattress with cover – 2 - Skin temperature probes – minimum 2 - Oxygen tubing/connector – 1 set - Rechargeable battery (installed) - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site delivery and functional testing - Verification of temperature accuracy and battery runtime - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site clinical training for NICU nurses and biomedical engineers - Includes operation, cleaning, maintenance, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance included during warranty period - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-19 (infant incubators) - CE MDR Regulation (EU) 2017/745 and FDA 510(k) certification - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/